20345716 | 2024 | XIENCE SKYPOINTÃ‚Â¿ | Unknown | ['Device Dislodged or Dislocated', 'No Clinical Signs, Symptoms or Conditions'] | THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE PRODUCTION RECORDS AND THE CORRECTIVE AND PREVENTIVE ACTIONS (CAPA) DATABASE WAS PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED NO INDICATION OF A LOT SPECIFIC ISSUE. THE INVESTIGATION DETERMINED THE REPORTED DIFFICULTIES AND SUBSEQUENT TREATMENT APPEAR TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. IT WAS REPORTED THAT DURING ADVANCEMENT, THE STENT DELIVERY SYSTEM (SDS) INTERACTED WITH THE PATIENTÂ¿S HEAVILY CALCIFIED, MILDLY TORTUOUS, AND 80% STENOSED LESION, RESULTING IN THE REPORTED STENT DISLODGEMENT. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT THE MID RIGHT CORONARY ARTERY (RCA) WITH HEAVY CALCIFICATION, MILD TORTUOSITY AND 80% STENOSIS. THE 3.5X33MM XIENCE SKYPOINT STENT DELIVERY SYSTEM (SDS) WAS ADVANCED; HOWEVER, THE STENT DISLODGED IN THE PROXIMAL RIGHT CORONARY ARTERY DUE TO THE CALCIFIED LESION. THE STENT WAS RETRIEVED WITH AN UNSPECIFIED BALLOON, AS THE STENT WAS CLOSE TO THE CATHETER. A NON-ABBOTT DEVICE WAS USED TO CONTINUE THE PROCEDURE. THERE WAS NO ADVERSE PATIENT SEQUELA. NO ADDITIONAL INFORMATION WAS PROVIDED.
20347506 | 2024 | XIENCE SIERRAÃ‚Â¿ | Male | ['Failure to Advance', 'Device Dislodged or Dislocated', 'Foreign Body In Patient'] | IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT A LESION IN AN UNSPECIFIED ARTERY. THE 4.0X18MM XIENCE SIERRA STENT DELIVERY SYSTEM (SDS) FAILED TO CROSS AND UPON REMOVAL, A DAMAGED AREA WAS NOTED TO HAVE NO MATERIAL. UNKNOWN IF THERE WAS INTERVENTION PERFORMED. THERE WAS NO CLINICALLY SIGNIFICANT DELAY IN THE PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED. B3: DATE OF EVENT ESTIMATED AS 09/01/2024. MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. A VISUAL INSPECTION WAS PERFORMED ON THE RETURNED DEVICE. THE REPORTED FAILURE TO ADVANCE COULD NOT BE TESTED AS IT WAS RELATED TO OPERATIONAL CONTEXT. THE REPORTED STENT DISLODGEMENT COULD NOT BE CONFIRMED. A REVIEW OF THE PRODUCTION RECORDS AND THE CORRECTIVE AND PREVENTIVE ACTIONS (CAPA) DATABASE WAS PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED NO INDICATION OF A LOT SPECIFIC ISSUE. THE INVESTIGATION DETERMINED THE REPORTED DIFFICULTIES AND PATIENT EFFECT APPEAR TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. IT WAS REPORTED THAT THE STENT DELIVERY SYSTEM (SDS) FAILED TO CROSS DUE TO INTERACTION WITH THE LESION. ADDITIONALLY, IT WAS REPORTED THAT THE STENT DISLODGED DURING USE. ANALYSIS OF THE RETURNED DEVICE WAS UNABLE TO CONFIRM THE STENT DISLODGEMENT, AS THE STENT WAS PRESENT ON THE SDS CRIMPED BETWEEN THE BALLOON MARKERS. ANALYSIS DID NOTE MATERIAL DEFORMATION ON THE STENT, IN THE FORM OF LIFTED STRUTS. THIS TYPE OF DAMAGE ALIGNS WITH INTERACTION WITH THE ANATOMY. IN THIS CASE, IT IS LIKELY THAT MANIPULATION OF THE SDS, WHEN RESISTANCE WITH THE ANATOMY WAS MET, CONTRIBUTED TO THE STENT DEFORMATION. ADDITIONALLY, IT WAS NOT REPORTED IF AN INTERVENTION WAS PERFORMED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING.
20357992 | 2024 | RESOLUTE ONYX RX | Male | ['Burst Container or Vessel', 'Leak/Splash', 'Deformation Due to Compressive Stress', 'Arrhythmia', 'Insufficient Information'] | AN ATTEMPT WAS MADE TO USE ONE RESOLUTE ONYX CORONARY DRUG ELUTING STENT TO TREAT A NON-TORTUOUS, NON-CALCIFIED LESION IN THE PROXIMAL RIGHT CORONARY ARTERY (RCA). THE DEVICE WAS NOT INSPECTED. NEGATIVE PREP WAS PERFORMED WITH NO ISSUES. THE LESION WAS PRE-DILATED. THE DEVICE DID PASS THROUGH A PREVIOUSLY DEPLOYED STENT. RESISTANCE WAS NOT ENCOUNTERED WHEN ADVANCING THE DEVICE. EXCESSIVE FORCE WAS NOT USED. IT WAS REPORTED THAT A BALLOON BURST/LEAK OCCURRED DURING STENT DEPLOYMENT. THE STENT CATHETER WAS PUT IN POSITION AND WHEN AN ATTEMPT WAS MADE TO INFLATE THE DEVICE NO PRESSURE SHOWED ON THE INFLATION DEVICE. THE INFLATION DEVICE WAS CONTINUED TO BE ROTATED BUT THERE WAS NO PRESSURE GOING INTO THE BALLOON, AND THE STENT WAS NOT EXPANDING. CONTRAST WAS SEEN JETTING BACK INTO THE AORTA AND IT WAS BELIEVED THAT THE BALLOON HAD RUPTURED. THE ISSUE OCCURRED ON THE FIRST INFLATION. A TOTAL OF 30CC OF CONTRAST WAS INJECTED TO EXPAND THE STENT ENOUGH TO BE ABLE TO REMOVE THE STENT DELIVERY SYSTEM AND GO BACK IN WITH AN NC BALLOON. THE PATIENT EXPERIENCED ST'S HOWEVER REMAINS STABLE. THE PATIENT IS ALIVE WITH NO FURTHER INJURY. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: THE STENT WAS SUCCESSFULLY IMPLANTED WITH THE NC BALLOON. THE PATIENT DID NOT REQUIRE ANY TREATMENT FOR THE STS EXPERIENCED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. PRODUCT ANALYSIS SUMMARY: THE DEVICE WAS RETURNED TO MEDTRONIC FOR EVALUATION. THE BALLOON RETURNED IN A POST INFLATED STATE AND KINKS WERE VISIBLE ON THE HYP OTUBE. BLOOD VISIBLE INSIDE THE BALLOON AND INFLATION LUMEN, INDICATING A LEAK. THE DEVICE FAILED NEGATIVE PREP. ON PRESSURISATION OF DEVICE A PIN HOLE LEAK OBSERVED, ON MOST PROXIMAL END OF THE WORKING LENGTH AND BALLOON FAILED TO MAINTAIN PRESSURE. BURST NOT CONFIRMED BUT LEAK CONFIRMED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20358105 | 2024 | ONYX FRONTIER | Male | ['Deflation Problem', 'Migration or Expulsion of Device', 'Device Damaged by Another Device', 'Device Dislodged or Dislocated', 'Obstruction/Occlusion', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. AN ATTEMPT WAS MADE TO USE ONE ONYX FRONTIER CORONARY DRUG ELUTING STENT TO TREAT A LESION. IT WAS REPORTED THAT BALLOON DEFLATION DIFFICULTIES OCCURRED, AND THE DEVICE WOULD NOT DEFLATE AT THE LESION SITE. A GUIDEWIRE WAS USED TO HELP DISLODGE THE BALLOON FROM THE STENT, HOWEVER, IN THE PROCESS THE STENT MOVED AND HAD TO BE RETRIEVED. THE PATIENT IS ALIVE WITH NO FURTHER INJURY. ADDITIONAL INFORMATION: ONE NC EUPHORA CORONARY BALLOON WAS ALSO BEING USED. THE ONYX FRONTIER STENT WAS IMPLANTED. IT WAS REPORTED THAT BALLOON DEFLATION DIFFICULTIES OCCURRED WITH THE NC EUPHORA BALLOON, AND THE DEVICE WOULD NOT DEFLATE AT THE LESION SITE. THE ONYX FRONTIER STENT WAS BEING POST-DILATED WITH THE NC EUPHORA BALLOON MID-STENT. ONE THE SECOND POST-DILATION INFLATION TO 6 ATM, THE BALLOON WASN'T FULLY DEFLATED. NEGATIVE WAS PULLED. THE STENT WAS FOUND IN THE ULNAR ARTERY AND WAS REMOVED BY A SNARE. THE NC EUPHORA BALLOON DEFLATED THE FIRST TIME BUT DID NOT DEFLATE AFTER THE SECOND INFLATION. THE CATHETER AND BALLOON DID NOT SEEM TO BE KINKED. THE PATIENT WAS FINE AND WAS DISCHARGED IN THE EVENING. PATIENT SEX AND GENDER PROVIDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: THE LESION BEING TREATED WAS IN THE LEFT ANTERIOR DESCENDING (LAD) ARTERY. THE NC EUPHORA BALLOON WAS INFLATED, AND IT APPEARED THAT THE STENT DEPLOYED APPROPRIATELY. WHEN THE GUIDEWIRE WAS REMOVED THE STENT WAS NOT VISIBLE UNDER FLUOROSCOPY. ADDITIONAL IMAGING WAS COMPLETED TO FIND THE STENT, BUT IT WAS NOT POSSIBLE TO LOCATE THE STENT. THE PATIENT RETURNED TO THE INPATIENT UNIT WHERE IT WAS IDENTIFIED THAT THE PATIENT'S HAND WAS COOL AND PULSELESS. THE PATIENT WAS SENT TO INTERVENTIONAL RADIOLOGY WHERE THE STENT WAS FOUND IN THE ULNAR ARTERY AND WAS REMOVED BY A SNARE. SECTION A. PATIENT INFORMATION: PATIENT DEMOGRAPHICS UPDATED. CORRECTION: ANNEX B CODE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20365162 | 2024 | SYNERGY | Female | ['Material Deformation', 'Vascular Dissection'] | IT WAS REPORTED THAT A DISSECTION OCCURRED. THE 90% STENOSED TARGET LESION WAS LOCATED IN THE MODERATELY TORTUOUS AND MODERATELY CALCIFIED LEFT CIRCUMFLEX ARTERY. THE LESION WAS PRE-DILATED WITH A 2.25 X 12 MM SEMI COMPLIANT BALLOON. A 2.50 X 48 MM SYNERGY DRUG-ELUTING STENT (DES) WAS ADVANCED, DEPLOYED AT 11 ATM AND POST DILATED WITH A 2.50 X 12 MM NON-COMPLIANT BALLOON AT 12 ATM. A TYPE A PROXIMAL EDGE DISSECTION AND STENT DAMAGE WAS NOTED. THE DISSECTION WAS COVERED BY DEPLOYING A 2.50 X 12 MM SYNERGY DES. THE PATIENT WAS STABLE AND DISCHARGED AFTER THE PROCEDURE.
20368647 | 2024 | RESOLUTE ONYX RX | Unknown | ['Deflation Problem', 'Inflation Problem', 'Leak/Splash', 'Difficult to Remove', 'Improper or Incorrect Procedure or Method', 'Activation, Positioning or Separation Problem', 'Material Deformation', 'Material Split, Cut or Torn', 'Insufficient Information'] | AN ATTEMPT WAS MADE TO USE ONE RESOLUTE ONYX CORONARY DRUG ELUTING STENT TO TREAT A LESION. THE STENT WAS USED OFF LABEL IN THE PULMONARY ARTERIES TO TREAT A SEVERELY SICK PEDIATRIC PATIENT. IT WAS REPORTED THAT INFLATION DIFFICULTIES OCCURRED DURING BALLOON INFLATION. THE BALLOON OF THE STENT COULD NOT BE INFLATED WITH PRESSURE. DUE TO THE LACK OF PRESSURE BUILD-UP, IT WAS ALSO NOT POSSIBLE TO COMPLETELY DEFLATE THE BALLOON AGAIN. AS A RESULT, THE BALLOON WAS SLIGHTLY FILLED AT THE DISTAL ENDS, BUT THE STENT COULD NOT BE EXPANDED, AS THE BALLOON WOULD END UP DEFLATING. AS THE STENT WAS IN THE PULMONARY ARTERIES IT WAS POSSIBLE TO PLACE THE STENT VIA THE VENTRICLE IN THE INFERIOR VENA CAVA (IVC). WITH A LOT OF FORCE, IT WAS POSSIBLE TO REMOVE THE STENT BALLOON VIA THE INSERTED SHEATH, HOWEVER, THE STENT COULD NOT BE REMOVED VIA THE SHEATH. THE STENT COULD NOT BE DEPLOYED IN THE IVC WITH THE STENT BALLOON. THE STENT HAD TO BE SECURED, AND IN THE THIRD ATTEMPT, THIS WAS ACHIEVED WITH DIFFICULTY WITH A HIGH-PRESSURE BALLOON. IT WAS SUSPECTED THAT THE STENT WAS TWISTED AND THEN WEDGED AND OPENED TOO QUICKLY UNDER A LOT OF PRESSURE. THE STENT WAS EXPANDED IN THE IVC. ON CLOSER INSPECTION OF THE SYSTEM IN VITRO, A LEAK WAS FOUND FROM WHICH THE CONTRAST AGENT/SODIUM CHLORIDE (NACL) MIXTURE WAS ESCAPING, SUGGESTING WHY IT WAS NOT POSSIBLE TO EXPAND OR DE-EXPAND THE BALLOON INTRAVASCULARLY. THE PATIENT DIED FOUR DAYS POST PROCEDURE. IT IS BELIEVED THAT THE DEATH WAS NOT MAINLY CAUSED BY THE STENT BALLOON MALFUNCTION, BUT THE MALFUNCTION DID AFFECT THE PROCEDURE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: INITIAL REPORTER NAME UPDATED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: AN ATTEMPT WAS MADE TO USE ONE RESOLUTE ONYX CORONARY DRUG ELUTING STENT TO TREAT CONGENITAL NARROWING OF THE PULMONARY ARTERY (PA). THERE WAS SEVERE STENOSIS IN THE PA. A 4X13MM NON-MEDTRONIC STENT WAS IMPLANTED IN THE RIGHT PA. IT WAS NOT DIFFICULT TO REMOVE THE PROTECTIVE SHEATH OR THE PACKAGING STYLETTE. THE DEVICE WAS INSPECTED BEFORE USE WITH NO ISSUES NOTED. THE DEVICE WAS NOT FLUSHED PRIOR TO USE. THE LESION WAS PRE-DILATED. RESISTANCE WAS NOT NOTED WHILE ADVANCING THE DEVICE TO THE LESION. THE DEVICE WAS NOT MOVED OR REPOSITIONED IN THE LESION WHILE INFLATED. THE ISSUE OCCURRED ON THE FIRST INFLATION AFTER AN INFLATION PRESSURE OF 16 ATM WAS APPLIED. A 0.25/0.75 CONCENTRATION OF CONTRAST/SALINE WAS USED. THE PHYSICIAN DID NOT ASSESS THAT THE DEATH EVENT WAS DIRECTLY RELATED TO THE USE OF THE RELEVANT DEVICE. THE CAUSE OF DEATH WAS MULTIMORBIDITY. THE PATIENT WAS NOT ON DUAL ANTIPLATELET THERAPY (DAPT) AT THE TIME OF THE EVENT. THERE WAS NO EVIDENCE OF THROMBUS. PRODUCT ANALYSIS: THE DEVICE WAS RETURNED FOR ANALYSIS. THERE ARE INDICATIONS THAT THE BALLOON HAD INFLATED/PARTIALLY INFLATED (BALLOON FOLDS EXPANDED). CRIMPS WERE VISIBLE ON THE EXPOSED BALLOON SURFACE; THE STENT DID NOT RETURN. TOOL/CLAMP MARKS WERE FOUND IN TWO LOCATIONS ON THE TRANSITION SHAFT. IT IS UNCLEAR WHAT CAUSED THESE CLAMP/TOOL MARKS AND/OR WHEN IT OCCURRED, HOWEVER WHATEVER CAUSED THE MARKS ALSO PUNCTURED THE WALL OF THE TRANSITION TUBING. WHEN ATTEMPTING TO RECREATE BALLOON INFLATION/DEVICE PRESSURIZATION, THREE LEAK POINTS WERE OBSERVED WHERE THE TRANSITION SHAFT WAS PUNCTURED. IT IS UNCLEAR AT THIS POINT IF THIS HAD ANY IMPACT ON THE INFLATION/DEFLATION ISSUES DURING THE PROCEDURE. NECKING WAS ALSO EVIDENT TO THE PROXIMAL BOND WHICH MAY HAVE CONTRIBUTED TO INFLATION/DEFLATION ISSUES. THE GUIDEWIRE ENTRY PORT WAS TORN. IMAGE ANALYSIS: A 4 SECOND VIDEO CLIP WAS PROVIDED FOR REVIEW. THE VIDEO APPEARED TO SHOW A LEAK ON PART OF THE DELIVERY SYSTEM OF THE DEVICE. IT IS NOT CLEAR WHAT AREA OF THE DEVICE THE LEAK IS COMING FROM. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. IMAGE ANALYSIS: THREE VIDEO CLIPS WERE PROVIDED TO SUPPORT THE INVESTIGATION. STILL SCREEN SHOTS HAVE BEEN TAKEN TO SUPPORT THE ANALYSIS OF THE IMAGES. IMAGES APPEAR TO SHOW THE ATTEMPTED PRESSURIZATION OF THE BALLOON OF A STENT DELIVERY SYSTEM, WITH THE DEVICE PRESENT WITHIN A SHEATH. THE BALLOON OF A DEVICE IS PRESSURIZED INSIDE THE SHEATH. THE PROXIMAL END OF THE BALLOON APPEARS TO HAVE EXPANDED BUT THE IMAGES DO NOT SHOW ANY EXPANSION ON THE DISTAL END OF THE BALLOON. BUT NO CONTRAST ESCAPE HAS BEEN IDENTIFIED TO CONFIRM A BALLOON BURST, SO, INFLATION DIFFICULTIES DO NOT APPEAR TO BE RELATED TO A BALLOON BURST. ANOTHER IMAGE APPEARS TO SHOW THE ATTEMPTED EXPANSION OF A STENT, ALSO INSIDE WHAT APPEARS TO BE A SHEATH. THE RESOLUTE ONYX STENT IS APPROVED FOR USE IN THE CORONARY VASCULATURE, THEREFORE USE OF THE DEVICE IN THE PEDIATRIC PULMONARY VASCULATURE IS OFF LABEL USE. THEREFORE, THE METHODOLOGY BEING ATTEMPTED CANNOT BE CONFIRMED. BUT THERE DOES APPEAR TO BE EVIDENCE THAT THE EXPANSION WITHIN A SHEATH HAS CONTRIBUTED TO INFLATION DIFFICULTIES. ADDITIONAL INFORMATION: THE STENT COULD NOT BE REMOVED VIA THE NON-MEDTRONIC 4F INTRODUCER SHEATH. CORRECTION PRODUCT ANALYSIS: NECKING WAS NOT EVIDENT TO THE PROXIMAL BOND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. CORRECTION: PRODUCT ANALYSIS: THE DEVICE WAS RETURNED FOR ANALYSIS. THERE ARE INDICATIONS ON THE RETURNED DEVICE THAT THE BALLOON HAD INFLATED/PARTIALLY INFLATED (BALLOON FOLDS EXPANDED). CRIMPS WERE VISIBLE ON THE EXPOSED BALLOON SURFACE; THE STENT WAS NOT RETURNED. MARKS WERE FOUND IN TWO LOCATIONS ON THE TRANSITION SHAFT. IT IS UNCLEAR WHAT CAUSED THESE MARKS AND/OR WHEN IT OCCURRED. THE WALL OF THE TRANSITION TUBING WAS ALSO PUNCTURED. WHEN ATTEMPTING TO PRESSURIZE THE DEVICE TO ENABLE INFLATION, THREE LEAK POINTS WERE OBSERVED WHERE THE TRANSITION SHAFT WAS PUNCTURED. IT IS UNCLEAR IF THE LEAKS IDENTIFIED HAD ANY IMPACT ON THE INFLATION/DEFLATION ISSUES DURING THE PROCEDURE. THE GUIDEWIRE ENTRY PORT WAS TORN. IMAGE ANALYSIS: THREE VIDEO CLIPS WERE PROVIDED TO SUPPORT THE INVESTIGATION. STILL SCREEN SHOTS HAVE BEEN TAKEN TO SUPPORT THE ANALYSIS OF THE IMAGES. IMAGES APPEAR TO SHOW THE ATTEMPTED PRESSURIZATION OF THE BALLOON OF A STENT DELIVERY SYSTEM. THE BALLOON OF A DEVICE IS PRESSURISED. THE PROXIMAL END OF THE BALLOON APPEARS TO HAVE EXPANDED BUT THE IMAGES DO NOT SHOW ANY EXPANSION ON THE DISTAL END OF THE BALLOON. NO CONTRAST ESCAPE HAS BEEN IDENTIFIED TO CONFIRM A BALLOON BURST, SO, INFLATION DIFFICULTIES DO NOT APPEAR TO BE RELATED TO A BALLOON BURST. ANOTHER IMAGE APPEARS TO SHOW THE ATTEMPTED EXPANSION OF A STENT. THE RESOLUTE ONYX STENT IS APPROVED FOR USE IN THE CORONARY VASCULATURE, THEREFORE USE OF THE DEVICE IN THE PAEDIATRIC PULMONARY VASCULATURE IS OFF LABEL USE. THEREFORE, THE METHODOLOGY BEING ATTEMPTED CANNOT BE CONFIRMED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20397282 | 2024 | XIENCE ALPINE | Female | ['Difficult to Advance', 'Vascular Dissection'] | MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. D4: A PARTIAL UDI WAS PROVIDED AS THE LOT NUMBER IS NOT KNOWN. IT WAS REPORTED THAT THE PROCEDURE WAS PERFORMED TO TREAT A LESION IN THE MIDDLE THIRD LEFT ANTERIOR DESCENDING (LAD) CORONARY ARTERY WITH MILD CALCIFICATION, MILD TORTUOSITY AND 80% STENOSIS. A 2.5X23MM XIENCE ALPINE STENT WAS IMPLANTED. POST DILATION WAS ATTEMPTED BY THE NC TREK NEO BALLOON DILATATION CATHETER (BDC) WHICH WAS ADVANCED OVER A 0.14 WHISPER GUIDE WIRE (GW). THE BDC WAS UNABLE TO PASS THROUGH THE STENT DUE TO SHAFT THICKNESS AND RESISTANCE WAS FELT CROSSING THE LESION. DURING BDC POSITIONING, THE GW AND NON-ABBOTT GUIDE CATHETER LOST ACCESS [DISPLACED], AND A DISSECTION WAS NOTED AT THE DISTAL EDGE OF THE STENT. A 2.25 X33 XIENCE ALPINE STENT WAS USED TO TREAT THE DISSECTION. ANOTHER NC BDC WAS USED TO CONTINUE THE PROCEDURE. THERE WERE NO REPORTED ADVERSE PATIENT SEQUELA. AN APPROXIMATE DELAY OF 50 MINUTES OCCURRED DUE THE DISPLACED DEVICES. NO ADDITIONAL INFORMATION WAS PROVIDED. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. THE REPORTED PATIENT EFFECT OF DISSECTION IS LISTED IN THE XIENCE ALPINE EVEROLIMUS ELUTING CORONARY STENT SYSTEM ELECTRONIC INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED ISSUE APPEARS TO BE RELATED TO OPERATIONAL CONTEXT. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE.
20417961 | 2024 | SYNERGY XD | Female | ['Material Rupture', 'Material Deformation', 'Activation Failure', 'No Clinical Signs, Symptoms or Conditions'] | IT WAS REPORTED THAT BALLOON RUPTURE AND STENT INSUFFICIENT APPOSITION OCCURRED. THE 80% STENOSED TARGET LESION WAS LOCATED IN THE NON-TORTUOUS AND MODERATELY CALCIFIED LEFT ANTERIOR DESCENDING ARTERY. A 2.50 X 48MM SYNERGY XD DRUG-ELUTING STENT WAS SELECTED FOR TREATMENT AND SUCCESSFULLY DELIVERED. HOWEVER, DURING THE INITIAL INFLATION USING A NON-BOSTON SCIENTIFIC INDEFLATOR AT FEW ATMOSPHERES FOR 20 TO 30 SECONDS, THE BALLOON WAS NOT HOLDING THE PRESSURE AND LEAKING CONTINUOUSLY. IT WAS POSSIBLE TO CONTINUE INFLATING UNTIL THE STENT MESH WAS EXPANDED ENOUGH FOR DEPLOYMENT. THE STENT DELIVERY SYSTEM WAS REMOVED, AND SHAFT KINK WAS NOTED. THE PROCEDURE WAS COMPLETED WITH ADDITIONAL DILATATIONS USING AN NC EMERGE BALLOON CATHETER TO FULLY EXPAND AND APPOSE THE STENT MESH WITHOUT ANY FURTHER ISSUES. NO PATIENT COMPLICATIONS WERE REPORTED.
20428678 | 2024 | SYNERGY XD | Unknown | ['Difficult to Remove', 'Device Dislodged or Dislocated', 'Foreign Body In Patient'] | IT WAS REPORTED THAT REMOVAL DIFFICULTIES WERE ENCOUNTERED, AND STENT DISLODGEMENT RESULTING IN ADDITIONAL INTERVENTION. THE 70% STENOSED TARGET LESION WAS LOCATED IN THE SEVERELY TORTUOUS AND MODERATELY CALCIFIED CORONARY ARTERY. A 2.75 X 16MM SYNERGY XD DRUG-ELUTING STENT WAS ADVANCED FOR TREATMENT. HOWEVER, DURING THE PROCEDURE, THE STENT WAS DISLODGED WHILE BEING PULLED BACK INTO THE GUIDE, TRYING TO RE-CROSS THE DEVICE BECOMING SNAGGED ON THE CA+ NODULE AND WAS STUCK AGAINST THE EDGE OF THE ROOFED GUIDE CATHETER. THE PROCEDURE WAS COMPLETED WITH THIS DEVICE WITHOUT ANY PATIENT COMPLICATIONS REPORTED.
20441096 | 2024 | SYNERGY XD | Unknown | ['Break', 'No Clinical Signs, Symptoms or Conditions'] | IT WAS REPORTED THAT SHAFT BREAK OCCURRED REQUIRING ADDITIONAL INTERVENTION. THE TARGET LESION, WHICH WAS 85% STENOSED, WAS LOCATED IN THE MILDLY TORTUOUS AND MODERATELY CALCIFIED LEFT MAIN ARTERY. A 4.00 X 48MM SYNERGY XD DRUG-ELUTING STENT WAS ADVANCED FOR DILATATION. HOWEVER, DURING THE REMOVAL OF THE DEVICE AFTER STENT DEPLOYMENT, THE SHAFT BROKE. AN ADDITIONAL GUIDEWIRE WAS PLACED, AND A BALLOON WAS UTILIZED TO PIN THE BROKEN DEVICE, SUCCESSFULLY EXTRACTING EVERYTHING FROM THE BODY. THE PROCEDURE WAS THEN COMPLETED USING AN ALTERNATE DEVICE, AND NO PATIENT COMPLICATIONS WERE REPORTED.
20442850 | 2024 | SYNERGY | Unknown | ['Material Deformation', 'Migration', 'Foreign Body In Patient'] | IT WAS REPORTED THAT STENT MIGRATION OCCURRED. THE 80% STENOSED TARGET LESION WAS LOCATED IN THE MODERATELY TORTUOUS AND MODERATELY CALCIFIED SAPHENOUS VEIN GRAFT (SVG). FOLLOWING PRE-DILATION WITH 3.00 AND 3.50 NC BALLOONS, A 4.00 X 20MM SYNERGY II DRUG ELUTING STENT WAS DEPLOYED TO TREAT THE TARGET LESION AND POST-DILATED WITH A 4.00 NC BALLOON. WHEN EVERYTHING WAS BEING REMOVED, THE STENT WAS DAMAGED AND CAME OUT OF THE SVG. THE PHYSICIAN TRIED TO REMOVE USING THE GUIDE, BUT THE STENT WAS LOST IN THE PATIENT. THE PROCEDURE WAS COMPLETED WITH ANOTHER STENT. THERE WERE NO PATIENT COMPLICATIONS. E1 INITIAL REPORTER ADDRESS 1: (B)(6). E1 INITIAL REPORTER ADDRESS 2: (B)(6).
20445504 | 2024 | PROMUS ELITE | Male | ['Fracture', 'No Clinical Signs, Symptoms or Conditions'] | IT WAS REPORTED THAT STENT FRACTURE OCCURRED REQUIRING ADDITIONAL INTERVENTION. THE 80% STENOSED TARGET LESION WAS LOCATED IN THE MODERATELY TORTUOUS AND MILDLY CALCIFIED LEFT ANTERIOR DESCENDING ARTERY. A 12 X 2.75 PROMUS ELITE DRUG-ELUTING STENT WAS SELECTED FOR TREATMENT. HOWEVER, DURING THE PROCEDURE, THE STENT FRACTURED IN THE MIDDLE. THE STENT WAS NOT REMOVED CONSIDERING THE SAFETY OF THE PATIENT, AND A NON-BOSTON SCIENTIFIC DEVICE WAS DEPLOYED TO COMPLETE THE PROCEDURE. NO PATIENT COMPLICATIONS WERE REPORTED.
20449489 | 2024 | RESOLUTE ONYX RX | Unknown | ['Device Dislodged or Dislocated', 'Insufficient Information', 'No Clinical Signs, Symptoms or Conditions'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. DURING A PROCEDURE ATTEMPTS WERE MADE TO USE TWO RESOLUTE ONYX CORONARY DRUG ELUTING STENTS TO TREAT A LESION LOCATED IN THE LEFT ANTERIOR DESCENDING (LAD) ARTERY. THE DEVICES WERE INSPECTED PRIOR TO USE WITH NO ISSUES NOTED. THE LESION WAS PRE-DILATED. THE RESOLUTE ONYX STENT (LOTÂ 0011927309) DID NOT PASS THROUGH A PREVIOUSLY DEPLOYED STENT. EXCESSIVE FORCE WAS NOT USED DURING DELIVERY OF BOTH DEVICES. IT WAS REPORTED THAT BOTH STENTS DISLODGED DURING DELIVERY THROUGH THE LESION. THE DISLODGED STENTS WERE REMOVED FROM THE PATIENT. WHEN THE RESOLUTE ONYX STENT (LOT NUMBERÂ 0226461484) WAS REMOVED FROM THE PATIENT, THE STENT STRUTS WERE NOTED TO BE DAMAGED. THE PATIENT WAS REPORTED TO BE ALIVE WITH NO INJURY. ADDITIONAL INFORMATION: IT WAS REPORTED THAT BOTH STENTS DISLODGED DURING ATTEMPTED DELIVERY THROUGH THE LESION. IT WAS STATED THAT WHEN THE FIRST STENT DISLODGED, IT WAS DECIDED TO USE ANOTHER SHORTER LENGTH, HOWEVER THIS STENT ALSO DISLODGED. RESISTANCE WAS NOTED DURING WITHDRAWAL OF THE DEVICES BUT EXCESSIVE FORCE WAS NOT USED. INTERVENTION WAS REQUIRED TO REMOVE THE DISLODGED STENTS FROM THE PATIENT AND A VASCULAR SURGEON WAS INVOLVED. THE PATIENT WAS REPORTED TO BE ALIVE WITH NO FURTHER INJURY. SECTION B1 ADVERSE EVENT/PRODUCT PROBLEM. H1 TYPE OF REPORTABLE EVENT. INITIAL REPORTER NAME AND PHONE NUMBER PROVIDED. CORRECTION: ANNEX D CODE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. PRODUCT ANALYSIS SUMMARY: THE DEVICE WAS RETURNED TO MEDTRONIC FOR EVALUATION. THE STENT WAS POSITIONED ON THE BALLOON BETWEEN THE MARKER BANDS AS PER POSITION SPECIFICATION, BUT DUE TO MID STENT DEFORMATION THE STENT DID NOT MEET VISUAL ACCEPTANCE SPECIFICATION. NO DEFORMATION WAS EVIDENT TO THE DISTAL TIP. THE INNER LUMEN PATENCY WAS VERIFIED WITH A 0.015 INCH MANDREL. NO OTHER DAMAGE EVIDENT TO THE REMAINDER OF THE DEVICE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: ANNEX D CODE ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20449495 | 2024 | RESOLUTE ONYX RX | Unknown | ['Device Dislodged or Dislocated', 'Material Deformation', 'Insufficient Information', 'No Clinical Signs, Symptoms or Conditions'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. DURING A PROCEDURE ATTEMPTS WERE MADE TO USE TWO RESOLUTE ONYX CORONARY DRUG ELUTING STENTS TO TREAT A LESION LOCATED IN THE LEFT ANTERIOR DESCENDING (LAD) ARTERY. THE DEVICES WERE INSPECTED PRIOR TO USE WITH NO ISSUES NOTED. THE LESION WAS PRE-DILATED. THE RESOLUTE ONYX STENT (LOTÂ 0011927309) DID NOT PASS THROUGH A PREVIOUSLY DEPLOYED STENT. EXCESSIVE FORCE WAS NOT USED DURING DELIVERY OF BOTH DEVICES. IT WAS REPORTED THAT BOTH STENTS DISLODGED DURING DELIVERY THROUGH THE LESION. THE DISLODGED STENTS WERE REMOVED FROM THE PATIENT. WHEN THE RESOLUTE ONYX STENT (LOT NUMBERÂ 0226461484) WAS REMOVED FROM THE PATIENT, THE STENT STRUTS WERE NOTED TO BE DAMAGED. THE PATIENT WAS REPORTED TO BE ALIVE WITH NO INJURY. ADDITIONAL INFORMATION: IT WAS REPORTED THAT BOTH STENTS DISLODGED DURING ATTEMPTED DELIVERY THROUGH THE LESION. IT WAS STATED THAT WHEN THE FIRST STENT DISLODGED, IT WAS DECIDED TO USE ANOTHER SHORTER LENGTH, HOWEVER THIS STENT ALSO DISLODGED. RESISTANCE WAS NOTED DURING WITHDRAWAL OF THE DEVICES BUT EXCESSIVE FORCE WAS NOT USED. INTERVENTION WAS REQUIRED TO REMOVE THE DISLODGED STENTS FROM THE PATIENT AND A VASCULAR SURGEON WAS INVOLVED. THE PATIENT WAS REPORTED TO BE ALIVE WITH NO FURTHER INJURY. SECTION B1 ADVERSE EVENT/PRODUCT PROBLEM. H1 TYPE OF REPORTABLE EVENT. INITIAL REPORTER NAME AND PHONE NUMBER PROVIDED. LOT NUMBER UPDATED. CORRECTION: ANNEX D CODE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. PRODUCT ANALYSIS SUMMARY: THE DEVICE WAS RETURNED TO MEDTRONIC FOR EVALUATION. THE STENT WAS NOT PRESENT ON THE BALLOON AND DID NOT RETURN FOR ANALYSIS. THE BALLOON FOLDS REMAINED INTACT. CRIMP IMPRESSIONS WERE VISIBLE ON THE EXPOSED BALLOON SURFACE. NO OTHER DAMAGE EVIDENT TO THE REMAINDER OF THE DEVICE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: ANNEX D CODES ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20460035 | 2024 | ONYX FRONTIER | Male | ['Device Dislodged or Dislocated', 'Positioning Problem', 'Insufficient Information', 'No Clinical Signs, Symptoms or Conditions'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. AN ATTEMPT WAS MADE TO USE ONE ONYX FRONTIER CORONARY DRUG ELUTING STENT TO TREAT A LESION. THE DEVICE WAS INSPECTED. NEGATIVE PREP WAS NOT PERFORMED. THE LESION WAS PRE-DILATED. THE DEVICE DID PASS THROUGH A PREVIOUSLY DEPLOYED STENT. RESISTANCE WAS ENCOUNTERED WHEN ADVANCING THE DEVICE. EXCESSIVE FORCE WAS NOT USED. IT WAS REPORTED THAT THE STENT WOULD NOT CROSS INTO THE AREA NEEDED FOR DEPLOYMENT. THE STENT WAS REMOVED AND UPON INSPECTION IT WAS NOTICED THAT THE STENT WAS NOT FULLY SEATED OVER THE BALLOON EVENLY BETWEEN THE MARKERS, AS USUAL. THE STENT APPEARED TO BE OFF BY A MILLIMETER OR TWO. A NEW SAME SIZE STENT WAS ABLE TO CROSS AND BE DEPLOYED. NO PATIENT INJURY REPORTED. ADDITIONAL INFORMATION: THE LESION BEING TREATED WAS A MODERATELY CALCIFIED, MILDLY TORTUOUS LESION WITH 70 TO 80% STENOSIS IN THE MID LEFT ANTERIOR DESCENDING ARTERY (LAD). THE DEVICE WAS INSPECTED WITH NO ISSUES NOTED. RESISTANCE WAS NOT NOTED DURING THE WITHDRAWAL OF THE DEVICE. A NEW SAME SIZE ONYX FRONTIER STENT WAS ABLE TO CROSS AND BE DEPLOYED. THE PATIENT WAS TRANSFERRED TO THE ICU AND EXPIRED TWO DAYS LATER. PATIENT HISTORY PROVIDED. PATIENT AGE SEX, GENDER, WEIGHT PROVIDED. B1 ADVERSE EVENT/PRODUCT PROBLEM. B2. OUTCOME ATTRIBUTED TO ADVERSE EVENT. H1. TYPE OF REPORTABLE EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. PRODUCT ANALYSIS: THE DEVICE WAS RETURNED FOR ANALYSIS. THERE WAS NO DAMAGE TO THE STENT AND THE STENT WAS POSITIONED BETWEEN THE TWO MARKER BANDS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20462580 | 2024 | SYNERGY MEGATRON | Unknown | ['Inflation Problem', 'Material Deformation', 'Failure to Deflate', 'No Clinical Signs, Symptoms or Conditions'] | IT WAS REPORTED THAT A BALLOON ISSUE OCCURRED. THE PATIENT PRESENTED FOR A COMPLEX PERCUTANEOUS CORONARY INTERVENTION. THE 4.50 X 28MM SYNERGY MEGATRON DRUG ELUTING STENT WAS ADVANCED TO TREAT THE TARGET LESION. DURING DEPLOYMENT, IT WAS NOTED THAT THE BALLOON WAS ONLY PARTIALLY INFLATED AND WAS UNABLE TO BE DEFLATED OR INFLATED FURTHER. THE STENT WAS SNARED AND THE ENTIRE SYSTEM REMOVED. THE PROCEDURE WAS COMPLETED USING AN ALTERNATE DEVICE. THERE WERE NO PATIENT COMPLICATIONS. THE DEVICE WAS NOT RETURNED FOR ANALYSIS; THEREFORE, A TECHNICAL ANALYSIS COULD NOT BE PERFORMED. HOWEVER, IMAGES PROVIDED BY THE CUSTOMER SHOW THE DEVICE LOADED INTO THE GUIDE CATHETER AND CONFIRM THE DAMAGE TO THE STENT. IT WAS REPORTED THAT A BALLOON ISSUE OCCURRED. THE PATIENT PRESENTED FOR A COMPLEX PERCUTANEOUS CORONARY INTERVENTION. THE 4.50 X 28MM SYNERGY MEGATRON DRUG ELUTING STENT WAS ADVANCED TO TREAT THE TARGET LESION. DURING DEPLOYMENT, IT WAS NOTED THAT THE BALLOON WAS ONLY PARTIALLY INFLATED AND WAS UNABLE TO BE DEFLATED OR INFLATED FURTHER. THE STENT WAS SNARED AND THE ENTIRE SYSTEM REMOVED. THE PROCEDURE WAS COMPLETED USING AN ALTERNATE DEVICE. THERE WERE NO PATIENT COMPLICATIONS.
20468078 | 2024 | ORSIRO MISSION (US) 3.5/15 | Unknown | ['Patient Device Interaction Problem', 'Chest Pain', 'Thrombosis/Thrombus'] | COMBINATION PRODUCT: YES. AN ORSIRO MISSION DRUG-ELUTING STENT SYSTEM WAS SELECTED FOR TREATMENT OF THE MID CX. IMMEDIATELY AFTER PROCEDURE AN IN STENT THROMBOSIS OCCURRED. PATIENT DEVELOPED CHEST PAIN. RE-BALLOONING OF THE STENT WAS PERFORMED AND ADDITIONAL HEPARIN WAS GIVEN. COMBINATION PRODUCT: YES, THE AFFECTED PRODUCT WAS NOT RETURNED TO BIOTRONIK AND COULD THEREFORE NOT BE SUBJECTED TO A TECHNICAL INVESTIGATION. IMAGES OF THE CASE SUCH AS ANGIOGRAPHIES OR OTHER RELEVANT CLINICAL INFORMATION COULD NOT BE OBTAINED. REVIEW OF THE PRODUCTION DOCUMENTATION CONFIRMED THAT THE DEVICE IN QUESTION WAS MANUFACTURED ACCORDING TO SPECIFICATIONS AND FULFILLED ALL THE REQUIREMENTS OF IN?PROCESS AND FINAL INSPECTION. BASED ON THE REVIEWED DOCUMENTATION NO DEVICE DEFICIENCY OR MANUFACTURING-RELATED ROOT CAUSE COULD BE IDENTIFIED. SINCE MEDICATION WAS CHANGED AFTER THE ADVERSE EVENT, DAPT EFFECTIVENESS MAY HAVE BEEN A CONTRIBUTING FACTOR TO THE REPORTED EVENT WHICH COULD NOT BE CONFIRMED. PLEASE NOTE THAT, ACCORDING TO THE ORSIRO MISSION US IFU, ADMINISTRATION OF APPROPRIATE ANTICOAGULANT, ANTIPLATELET AND VASODILATION THERAPY IS CRITICAL TO SUCCESSFUL STENT IMPLANTATION.
20468091 | 2024 | ORSIRO MISSION (US) 3.5/26 | Unknown | ['Patient Device Interaction Problem', 'Chest Pain', 'Thrombosis/Thrombus'] | COMBINATION PRODUCT: YES. AN ORSIRO MISSION DRUG-ELUTING STENT SYSTEM WAS SELECTED FOR TREATMENT OF THE PROXIMAL LAD. IMMEDIATELY AFTER PROCEDURE AN IN-STENT THROMBOSIS OCCURRED. PATIENT DEVELOPED CHEST PAIN. RE-BALLOONING OF THE STENT WAS PERFORMED AND ADDITIONAL HEPARIN WAS GIVEN. PATIENT WAS TREATED FOR HEPARIN-INDUCED THROMBOCYTOPENIA (HIT). PHYSICIAN SHARED THAT DURING THE CASE THE PATIENTS ACT WOULD BE 182, GIVE HEPARIN WOULD GO TO 350 THEN WHEN RE-CHECKED AGAIN WOULD BE BACK DOWN TO 180. COMBINATION PRODUCT: YES. THE AFFECTED PRODUCT WAS NOT RETURNED TO BIOTRONIK AND COULD THEREFORE NOT BE SUBJECTED TO A TECHNICAL INVESTIGATION. IMAGES OF THE CASE SUCH AS ANGIOGRAPHIES OR OTHER RELEVANT CLINICAL INFORMATION COULD NOT BE OBTAINED. REVIEW OF THE PRODUCTION DOCUMENTATION CONFIRMED THAT THE DEVICE IN QUESTION WAS MANUFACTURED ACCORDING TO SPECIFICATIONS AND FULFILLED ALL THE REQUIREMENTS OF IN?PROCESS AND FINAL INSPECTION. BASED ON THE REVIEWED DOCUMENTATION NO DEVICE DEFICIENCY OR MANUFACTURING-RELATED ROOT CAUSE COULD BE IDENTIFIED. CONSIDERING THE HIT-SITUATION OF THE PATIENT DESCRIBED BY THE PHYSICIAN, THE MOST PROBABLE ROOT CAUSE FOR THE COMPLAINT EVENT IS THE PATIENTS ANATOMY. PLEASE NOTE THAT, ACCORDING TO THE ORSIRO MISSION US IFU, ADMINISTRATION OF APPROPRIATE ANTICOAGULANT, ANTIPLATELET AND VASODILATION THERAPY IS CRITICAL TO SUCCESSFUL STENT IMPLANTATION.
20495643 | 2024 | SYNERGY XD | Unknown | ['Failure to Advance', 'Device Dislodged or Dislocated', 'No Clinical Signs, Symptoms or Conditions'] | IT WAS REPORTED THAT STENT DISLODGEMENT INSIDE THE PATIENT OCCURRED. THE PATIENT PRESENTED WITH CORONARY ARTERY DISEASES AND UNDERWENT PERCUTANEOUS CORONARY INTERVENTION. THE 80% STENOSED TARGET LESION WAS LOCATED IN THE NON-TORTUOUS AND MODERATELY CALCIFIED LEFT ANTERIOR DESCENDING ARTERY. A 3.00 X 16MM SYNERGY XD DRUG-ELUTING STENT WAS ADVANCED FOR TREATMENT BUT FAILED TO CROSS THE LESION. HOWEVER, WHILE REMOVING THE STENT DELIVERY CATHETER OUT OF THE BODY, THE STENT DISLODGED IN THE LEFT MAIN WITHOUT THE INVOLVEMENT OF INFLATION DEVICE. A SNARE WAS THEN USED TO RETRIEVE THE STENT COMPLETELY, AND A NON-BOSTON SCIENTIFIC DEVICE WAS USED TO COMPLETE THE PROCEDURE SUCCESSFULLY. NO PATIENT COMPLICATIONS WERE REPORTED.
20499920 | 2024 | SYNERGY XD | Unknown | ['Break', 'Difficult to Remove', 'Failure to Deflate', 'Foreign Body In Patient', 'Vascular Dissection', 'Embolism/Embolus', 'Reduced Blood Flow'] | IT WAS REPORTED THAT REMOVAL DIFFICULTIES OCCURRED. THE TARGET LESION WAS LOCATED IN THE PROXIMAL FIRST OBTUSE MARGINAL (OM). A 2.75 X 38MM SYNERGY XD DRUG-ELUTING STENT WAS ADVANCED FOR TREATMENT AND DEPLOYED WITH NO ISSUES. HOWEVER, WHEN THE STENT DELIVERY SYSTEM (SDS) BALLOON WAS FULLY DEFLATED, IT REMAINED STUCK IN THE STENT. MULTIPLE PULLS AND MULTIPLE INFLATIONS AND DEFLATIONS ENSUED WITHOUT SUCCESS, AND A FAILED ATTEMPT WAS MADE TO TRY AND GET A SECOND WIRE PAST THE STENT. A HARD PULL WAS MADE TO THE SDS CATHETER AT WHICH TIME THE SHAFT BROKE. THE PROXIMAL HALF WAS RETRIEVED, BUT THE DISTAL HALF REMAINED STUCK IN THE STENT. THE CATHETER TIP EMBOLIZED IN THE VESSEL AND THE DOCTOR OPTED TO JAIL THE TIP INTO THE OM WITH ANOTHER STENT, EFFECTIVELY SHUTTING DOWN THE OM. THE BALLOON BEING STUCK AND THE FORCED NEEDED TO REMOVE IT CAUSED A NON-FLOW LIMITING DISSECTION IN THE OSTIUM OF THE OM. THE PATIENT WAS ADMITTED TO THE HOSPITAL BEYOND STANDARD OF CARE AND TREATED WITH HEAVY ANTICOAGULATION AND OBSERVATION. THE PATIENT WAS STABLE FOLLOWING THE PROCEDURE AND LATER DISCHARGED. DEVICE EVALUATED BY MFR.: SYNERGY XD MR US 2.75 X 38MM WAS RETURNED FOR ANALYSIS. A VISUAL AND TACTILE AND MICROSCOPIC EXAMINATION IDENTIFIED NO KINKS OR DAMAGES ALONG THE ENTIRE LENGTH OF THE HYPOTUBE. A VISUAL AND TACTILE EXAMINATION OF THE DISTAL EXTRUSION IDENTIFIED A BREAK IN THE MIDSHAFT AT 8CM FROM THE END TAB OF THE LASER CUT SECTION. THE DISTAL SECTION OF THE BREAK INCLUDING THE INNER AND OUTER DISTAL EXTRUSIONS, THE BALLOON AND TIP SECTION OF THE DEVICE WAS NOT RETURNED FOR ANALYSIS. MICROSCOPIC INSPECTION OF THE BREAK SITE REVEALED THE MIDSHAFT EXTRUSION WAS STRETCHED. IT WAS REPORTED THAT REMOVAL DIFFICULTIES OCCURRED. THE TARGET LESION WAS LOCATED IN THE PROXIMAL FIRST OBTUSE MARGINAL (OM). A 2.75 X 38MM SYNERGY XD DRUG-ELUTING STENT WAS ADVANCED FOR TREATMENT AND DEPLOYED WITH NO ISSUES. HOWEVER, WHEN THE STENT DELIVERY SYSTEM (SDS) BALLOON WAS FULLY DEFLATED, IT REMAINED STUCK IN THE STENT. MULTIPLE PULLS AND MULTIPLE INFLATIONS AND DEFLATIONS ENSUED WITHOUT SUCCESS, AND A FAILED ATTEMPT WAS MADE TO TRY AND GET A SECOND WIRE PAST THE STENT. A HARD PULL WAS MADE TO THE SDS CATHETER AT WHICH TIME THE SHAFT BROKE. THE PROXIMAL HALF WAS RETRIEVED, BUT THE DISTAL HALF REMAINED STUCK IN THE STENT. THE CATHETER TIP EMBOLIZED IN THE VESSEL AND THE DOCTOR OPTED TO JAIL THE TIP INTO THE OM WITH ANOTHER STENT, EFFECTIVELY SHUTTING DOWN THE OM. THE BALLOON BEING STUCK AND THE FORCED NEEDED TO REMOVE IT CAUSED A NON-FLOW LIMITING DISSECTION IN THE OSTIUM OF THE OM. THE PATIENT WAS ADMITTED TO THE HOSPITAL BEYOND STANDARD OF CARE AND TREATED WITH HEAVY ANTICOAGULATION AND OBSERVATION. THE PATIENT WAS STABLE FOLLOWING THE PROCEDURE AND LATER DISCHARGED.
20500367 | 2024 | XIENCE SKYPOINTÃ‚Â¿ | Unknown | ['Failure to Advance', 'Device Dislodged or Dislocated', 'Device Embedded In Tissue or Plaque'] | IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT A LESION IN THE CIRCUMFLEX (CX) ARTERY WITH 90% STENOSIS AND MODERATE TORTUOSITY. THE 3.5X23MM XIENCE SKYPOINT STENT DELIVERY SYSTEM WAS BEING ADVANCED WITHOUT SUFFICIENT GUIDE SUPPORT AND FAILED TO CROSS DUE TO THE ANATOMY WHEN THE STENT DISLODGED IN THE LEFT MAIN ARTERY. AN ATTEMPT TO SNARE THE STENT WAS UNSUCCESSFUL. ANOTHER 3.5X28MM XIENCE SKYPOINT STENT WAS DEPLOYED TO CRUSH THE DISLODGED STENT AND COMPLETE THE PROCEDURE. THE PATIENT IS DOING WELL. THERE WAS NO CLINICALLY SIGNIFICANT DELAY IN THE PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED. MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. A VISUAL INSPECTION ANALYSIS WAS PERFORMED ON THE RETURNED DEVICE. DEVICE DISLODGE OR DISLOCATED AND FAILURE TO ADVANCE WAS CONFIRMED. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. THE INVESTIGATION DETERMINED THE REPORTED FAILURE TO ADVANCE, DISLODGED OR DISLOCATED STENT AND SUBSEQUENT TREATMENTS APPEAR TO BE RELATED TO CIRCUMSTANCES OF THE PROCEDURE. IT WAS REPORTED THAT THE STENT DELIVERY SYSTEM (SDS) MET RESISTANCE WITH THE LESION, RESULTING IN A FAILURE TO ADVANCE AND STENT DISLODGEMENT. IN THIS CASE, ANALYSIS OF THE RETURNED DEVICE CONFIRMED THE REPORTED STENT DISLODGEMENT. THIS TYPE OF DAMAGE LIKELY OCCURRED DUE TO THE INTERACTION WITH THE CHALLENGING LESION. ADDITIONALLY, IT WAS REPORTED THAT THE DISLODGED STENT WAS EMBEDDED IN THE PATIENT'S ANATOMY. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE.
20505472 | 2024 | ONYX FRONTIER | Male | ['Device Dislodged or Dislocated', 'Device Embedded In Tissue or Plaque', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. AN ATTEMPT WAS MADE TO USE ONE ONYX FRONTIER CORONARY DRUG-ELUTING STENT (DES) TO TREAT A MILDLY CALCIFIED, MILDLY TORTUOUS LESION IN THE PROXIMAL CIRCUMFLEX ARTERY (CX). THE DEVICE WAS INSPECTED WITH NO ISSUES NOTED. NEGATIVE PREP WAS PERFORMED WITH NO ISSUES NOTED. THE LESION WAS PRE-DILATED. THE DEVICE DID NOT PASS THROUGH A PREVIOUSLY DEPLOYED STENT. RESISTANCE WAS NOT ENCOUNTERED WHEN ADVANCING THE DEVICE. EXCESSIVE FORCE WAS NOT USED. IT WAS REPORTED THAT STENT DISLODGEMENT OCCURRED DURING DELIVERY TO/AT THE LESION. IT WAS DETAILED THAT THE STENT CAME OFF THE BALLOON IN THE ARTERY IN THE PROXIMAL CIRCUMFLEX. ANOTHER STENT WAS DEPLOYED TO CRUSH THE LOOSE STENT INTO THE ARTERIAL WALL. THE PATIENT WAS STABLE POST-PROCEDURE. THE PATIENT IS ALIVE WITH NO FURTHER INJURY. ADDITIONAL INFORMATION: RESISTANCE WAS NOT NOTED DURING THE WITHDRAWAL OF THE DEVICE. EXCESSIVE FORCE WAS NOT USED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20507179 | 2024 | XIENCE SKYPOINTÃ‚Â¿ | Male | ['Migration', 'No Clinical Signs, Symptoms or Conditions'] | MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. IT WAS REPORTED THAT THE PROCEDURE WAS PERFORMED TO TREAT A LESION IN THE LEFT MAIN CORONARY ARTERY WITH MILD CALCIFICATION AND WHERE A PREVIOUSLY IMPLANTED, UNSPECIFIED STENT, WAS LOCATED. THE 4.0X8MM XIENCE SKYPOINT STENT WAS DEPLOYED TO OVERLAP THE IMPLANTED STENT; HOWEVER, THE OVERLAP WAS MISSED BY 1MM. THE IMPLANTED STENT REMAINED WITHIN THE TARGET LESION AND NOT IN HEALTH TISSUE. POST-DILATION WAS PERFORMED WITH AN UNSPECIFIED 4.5X08 NON-COMPLIANT BALLOON WITHOUT ISSUE. INTRAVASCULAR ULTRASOUND (IVUS) WAS THEN USED, NOTING THAT THE STENT HAD MIGRATED OVER THE TOP OF THE GUIDING CATHETER AND INTO THE FEMORAL ARTERY WHERE IT REMAINED. AN UNSPECIFIED STENT WAS IMPLANTED IN THE TARGET LESION. ADDITIONAL HOSPITALIZATION OCCURRED FOR VASCULAR SURGERY TO REMOVE THE MIGRATED STENT BY CUT DOWN. THERE WERE NO ADVERSE PATIENT SEQUELA. NO ADDITIONAL INFORMATION WAS PROVIDED. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE INFORMATION PROVIDED, A DEFINITIVE CAUSE FOR THE REPORTED ISSUE COULD NOT BE DETERMINED. THE REPORTED TREATMENT OF SURGICAL INTERVENTION APPEARS TO BE RELATION TO OPERATIONAL CONTEXT OF THE PROCEDURE. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE.
20509695 | 2024 | XIENCE SIERRAÃ‚Â¿ | Unknown | ['Break', 'No Clinical Signs, Symptoms or Conditions'] | MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. IT WAS REPORTED THAT THE PROCEDURE WAS PERFORMED TO TREAT A LESION IN THE LEFT ANTERIOR DESCENDING ARTERY (LAD). A 3.25X28MM XIENCE SIERRA DRUG ELUTING STENT (DES) WAS IMPLANTED AT THE TARGET LESION; HOWEVER, THE IMPLANTED STENT WAS NOTED TO FRACTURE. A 3.0X38MM XIENCE SIERRA DES WAS DEPLOYED TO COVER THE FRACTURED STENT, AND THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE PATIENT EFFECTS AND NO CLINICALLY SIGNIFICANT DELAY IN THE PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE INFORMATION PROVIDED, A DEFINITIVE CAUSE FOR THE REPORTED ISSUE COULD NOT BE DETERMINED. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE.
20509805 | 2024 | XIENCE XPEDITION 48 EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Unknown | ['Break', 'No Clinical Signs, Symptoms or Conditions'] | IT WAS REPORTED THAT THE PROCEDURE WAS PERFORMED TO TREAT A LESION IN THE LEFT ANTERIOR DESCENDING ARTERY (LAD). A 2.5X48MM XIENCE XPEDITION DRUG ELUTING STENT (DES) WAS IMPLANTED AT THE TARGET LESION; HOWEVER, THE IMPLANTED STENT WAS NOTED TO FRACTURE. A 2.5X38MM XIENCE ALPINE DES WAS DEPLOYED TO COVER THE FRACTURED STENT, AND THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE PATIENT EFFECTS AND NO CLINICALLY SIGNIFICANT DELAY IN THE PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED. MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE INFORMATION PROVIDED, A DEFINITIVE CAUSE FOR THE REPORTED ISSUE COULD NOT BE DETERMINED. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. E1 - ACCOUNT NUMBER UPDATED FROM (B)(6).
20523534 | 2024 | SYNERGY XD | Unknown | ['Failure to Advance', 'Thrombosis/Thrombus'] | IT WAS REPORTED THAT THROMBOSIS OCCURRED. THE 99% STENOSED TARGET LESION WAS LOCATED IN THE NON-TORTUOUS AND HIGH-GRADE LESION OF THE PROXIMAL LEFT ANTERIOR DESCENDING ARTERY. A 2.50 X 16MM SYNERGY XD DRUG-ELUTING STENT WAS ADVANCED; BUT FAILED TO CROSS THE LESION AND THROMBOSIS OCCURRED. MULTIPLE BALLOON INFLATIONS AND AN ATTEMPT TO PLACE THE STENT WERE PERFORMED. THE PATIENT WAS PLACED ON AN INTRA-AORTIC BALLOON PUMP TO BE SAFELY TRANSFERRED TO ANOTHER FACILITY. THE PATIENT WAS IN STABLE CONDITION POST PROCEDURE.
20531556 | 2024 | ONYX FRONTIER | Unknown | ['Difficult to Remove', 'Device Dislodged or Dislocated', 'Material Deformation', 'Positioning Problem', 'Device Embedded In Tissue or Plaque', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. AN ATTEMPT WAS MADE TO USE ONE ONYX FRONTIER CORONARY DRUG ELUTING STENT (DES) TO TREAT A MODERATELY CALCIFIED LESION IN THE LEFT MAIN (LM) CORONARY ARTERY AND LEFT ANTERIOR DESCENDING (LAD) ARTERY. RESISTANCE WAS ENCOUNTERED WHEN ADVANCING THE DEVICE. IT WAS REPORTED THAT STENT DISLODGEMENT OCCURRED DURING REMOVAL FOLLOWING A FAILED DELIVERY. IT WAS DETAILED THAT THE LM WAS ENGAGED WITH AN EBU 3.75 GUIDE AND WAS WIRED WITH A NON-MEDTRONIC (MDT) WIRE WITH SOME DIFFICULTY. THE LESION WAS PRE-DILATED WITH A 2.75MM BALLOON. THE LESION WAS THEN STENTED WITH A 3.0X8MM DES AT 16 ATM. THERE WAS A PROXIMAL EDGE DISSECTION. THE ONYX FRONTIER STENT WAS BEING USED TO COVER THE DISSECTION, BUT THE STENT WAS UNABLE TO PASS DUE TO CALCIUM IN THE PROXIMAL LAD AND LM. A TELESCOPE GUIDE EXTENSION CATHETER (GEC) WAS USED TO ASSIST IN PASSING THE STENT, BUT THE STENT DISLODGED FROM THE BALLOON VERY EASILY, LIKELY DUE TO STENT MALFUNCTION, AND WAS FLOATING IN THE PROXIMAL LAD AND LM. AN ATTEMPT WAS MADE TO SNARE THE STENT WITH A 2.0MM BALLOON DEPLOYED WITHIN THE STENT TO TRY PULL IT INTO THE GUIDE, HOWEVER THE STENT WOULD NOT COME INTO THE GUIDE. A DECISION WAS MADE TO DEPLOY THE STENT. DUE TO TELESCOPE MALFUNCTION AND DIFFICULTY, ACCESS WITH THE GUIDE AND WIRE WAS LOST. THE LAD WAS RE-WIRED WITH A NON-MDT WIRE, AND AN INTRAVASCULAR ULTRASOUND (IVUS) WAS PERFORMED WHICH SHOWED A PARTIALLY EXPANDED STENT IN THE LAD AND AN UNDER-EXPANDED DISLODGED STENT IN THE LM. THE DISLODGED STENT WAS CRUSHED AGAINST LM VESSEL WALL. THE LAD AND LM WERE DILATED WITH A 3.5 BALLOON AND THE LESION WAS THEN STENTED WITH A 4.0X12MM DRUG ELUTING STENT IN THE LM AND LAD AT 16 ATM. IVUS WAS PERFORMED AGAIN WHICH SHOWED A WELL-EXPANDED STENT AND A SUCCESSFULLY CRUSHED STENT IN THE LM. THE LAD WAS FURTHER DILATED WITH A 3.0MM BALLOON, AND A 3.0X38MM STENT WAS PLACED OVER THE INITIAL DISSECTION WITH THE HELP OF A NON-MDT GEC AND DEPLOYED AT 16 ATM. IVUS WAS PERFORMED WHICH SHOWED A WELL EXPANDED STENT. THE PATIENT IS ALIVE WITH NO FURTHER INJURY. PRODUCT ANALYSIS: THE DEVICE WAS RETURNED FOR ANALYSIS. THE DEVICE WAS KINKED ON RETURN. THE STENT WAS NOT PRESENT ON THE BALLOON AND DID NOT RETURN FOR ANALYSIS. THE BALLOON FOLDS REMAINED INTACT. CRIMP IMPRESSIONS WERE VISIBLE ON THE EXPOSED BALLOON SURFACE. NO OTHER DAMAGE EVIDENT TO THE REMAINDER OF THE DEVICE. ADDITIONAL INFORMATION: DEVICE LOT NUMBER AND DETAILS UPDATED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. THE LESION WAS INITIALLY STENTED WITH A 3.0X8MM ONYX FRONTIER DES AT 16 ATM. AFTER WHICH THERE WAS A PROXIMAL EDGE DISSECTION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20545256 | 2024 | SYNERGY XD | Unknown | ['Break', 'Device Dislodged or Dislocated', 'Arrhythmia', 'Obstruction/Occlusion'] | IT WAS REPORTED THAT STENT DISLODGEMENT OCCURRED RESULTING IN VESSEL OCCLUSION. THE PATIENT PRESENTED WITH A NON-ST-ELEVATION MYOCARDIAL INFARCTION (NSTEMI). A 3.50 X 20 MM SYNERGY XD DRUG-ELUTING STENT WAS TO BE DEPLOYED IN THE PROXIMAL CIRCUMFLEX. DURING THE PROCEDURE, THE PATIENT WENT INTO VENTRICULAR FIBRILLATION (VFIB) AND WAS BROUGHT BACK TO SINUS RHYTHM WHILE THE BALLOON WAS DEFLATED. UPON RETURNING TO FLUOROSCOPY, THE PHYSICIAN COULD STILL SEE THE MARKER BANDS OF THE BALLOON. HOWEVER, THE BALLOON CAME OFF THE SHAFT WHEN IT WAS PULLED OUT. THE PHYSICIAN THEN WAS UNABLE TO SNARE IT OR CROSS ANOTHER DEVICE, WHICH CAUSED THE VESSEL TO OCCLUDE. THE CIRCUMFLEX SHUT DOWN, RESULTING IN NO FLOW. THE PATIENT HAS BEEN INTUBATED AND IS IN THE INTENSIVE CARE UNIT (ICU). NO FURTHER ISSUES WERE REPORTED. B5. DESCRIBE EVENT OR PROBLEM UPDATED. H6. PATIENT, IMPACT AND DEVICE CODE UPDATED. IT WAS REPORTED THAT STENT DISLODGEMENT OCCURRED RESULTING IN VESSEL OCCLUSION. THE PATIENT PRESENTED WITH A NON-ST-ELEVATION MYOCARDIAL INFARCTION (NSTEMI). A 3.50 X 20 MM SYNERGY XD DRUG-ELUTING STENT WAS TO BE DEPLOYED IN THE PROXIMAL CIRCUMFLEX. DURING THE PROCEDURE, THE PATIENT WENT INTO VENTRICULAR FIBRILLATION (VFIB) AND WAS BROUGHT BACK TO SINUS RHYTHM WHILE THE BALLOON WAS DEFLATED. UPON RETURNING TO FLUOROSCOPY, THE PHYSICIAN COULD STILL SEE THE MARKER BANDS OF THE BALLOON. HOWEVER, THE BALLOON CAME OFF THE SHAFT WHEN IT WAS PULLED OUT. THE PHYSICIAN THEN WAS UNABLE TO SNARE IT OR CROSS ANOTHER DEVICE, WHICH CAUSED THE VESSEL TO OCCLUDE. THE CIRCUMFLEX SHUT DOWN, RESULTING IN NO FLOW. THE PATIENT HAS BEEN INTUBATED AND IS IN THE INTENSIVE CARE UNIT (ICU). NO FURTHER ISSUES WERE REPORTED. IT WAS FURTHER REPORTED THAT THE STENT WAS NOT DETACHED AS WHAT WAS PREVIOUSLY REPORTED, RATHER IT WAS A SHAFT BREAK THAT LEFT THE BALLOON SECTION BEHIND.
20559794 | 2024 | PROMUS ELITE | Male | ['Fracture', 'No Clinical Signs, Symptoms or Conditions'] | E1: INITIAL REPORTER PHONE: (B)(6). IT WAS REPORTED THAT STENT FRACTURE OCCURRED. THE PATIENT WAS BEING TREATED WITH PERCUTANEOUS CORONARY INTERVENTION (PCI) OF A TORTUOUS ARTERY. FOLLOWING DEPLOYMENT OF A 28 X 2.75 PROMUS ELITE DRUG ELUTING STENT, THE STENT WAS NOTED TO BE FRACTURED. TWO ADDITIONAL STENTS WERE DEPLOYED TO COVER THE FRACTURE. THERE WERE NO PATIENT COMPLICATIONS REPORTED. E1: INITIAL REPORTER PHONE: (B)(6). IT WAS REPORTED THAT STENT FRACTURE OCCURRED. THE PATIENT WAS BEING TREATED WITH PERCUTANEOUS CORONARY INTERVENTION (PCI) OF A TORTUOUS ARTERY. FOLLOWING DEPLOYMENT OF A 28 X 2.75 PROMUS ELITE DRUG ELUTING STENT, THE STENT WAS NOTED TO BE FRACTURED. TWO ADDITIONAL STENTS WERE DEPLOYED TO COVER THE FRACTURE. THERE WERE NO PATIENT COMPLICATIONS REPORTED. IT WAS FURTHER REPORTED THAT THE STENT WAS NOT BROKEN INTO TWO PIECES BUT THAT TWO OF THE STRUTS WERE BROKEN.
20560174 | 2024 | PROMUS ELEMENT | Male | ['Activation Failure', 'Restenosis'] | AGENT JAPAN SV. IT WAS REPORTED THAT IN-STENT RESTENOSIS AND UNDER STENT EXPANSION OCCURRED. ON (B)(6) 2010, 90 % STENOSIS AT THE MID LEFT ANTERIOR DESCENDING CORONARY ARTERY (LAD) WAS TREATED WITH A 2.5 MM X 28 MM PROMUS STENT. ON (B)(6) 2021, THE SUBJECT PRESENTED WITH STABLE ANGINA AND WAS REFERRED FOR THE AGENT JAPAN SV STUDY. HEPARIN OR ANTITHROMBIN WAS ADMINISTERED AT THE TIME OF INDEX PROCEDURE. THE TARGET LESION WAS LOCATED AT THE MID LAD WITH 90% STENOSIS AND WAS 28 MM LONG LESION WITH A REFERENCE VESSEL DIAMETER WAS 2.5 MM. THE TARGET LESION WAS PREDILATED WITH A 2.50 MM X 15 MM BALLOON. FOLLOWING PRE-DILATION, THE TARGET LESION WAS TREATED WITH A 2.50 MM X 30 MM AGENT BALLOON, RESULTING IN 0% RESIDUAL STENOSIS AND TIMI FLOW 3. POST DILATION WAS NOT PERFORMED. INTRAVASCULAR ULTRASOUND (IVUS) REVEALED INSUFFICIENT STENT EXPANSION AT THE MID LAD. A NON-TARGET LESION LOCATED FROM THE DISTAL LEFT CIRCUMFLEX CORONARY ARTERY (LCX) WAS TREATED SUCCESSFULLY WITH BALLOON ANGIOPLASTY USING A 2.50 MM X 20 MM NON-BOSTON SCIENTIFIC (NON-BSC) BALLOON FOLLOWED BY A 2.25 MM X 20 MM SYNERGY STENT PRIOR TO THE TARGET LESION TREATMENT. THE SUBJECT WAS DISCHARGED ON ASPIRIN AND CLOPIDOGREL.
20567234 | 2024 | ONYX FRONTIER | Unknown | ['Device Dislodged or Dislocated', 'Positioning Problem', 'Device Embedded In Tissue or Plaque', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. AN ATTEMPT WAS MADE TO USE AN ONYX FRONTIER CORONARY DRUG ELUTING STENT TO TREAT A NON TORTUOUS, SEVERELY CALCIFIED LESION IN THE MID RIGHT CORONARY ARTERY (RCA). THE DEVICE WAS INSPECTED WITH NO ISSUES NOTED. NEGATIVE PREP WAS PERFORMED WITH NO ISSUES NOTED. THE LESION WAS PRE DILATED. THE DEVICE DID NOT PASS THROUGH A PREVIOUSLY-DEPLOYED STENT. RESISTANCE WAS ENCOUNTERED WHEN ADVANCING THE DEVICE AND EXCESSIVE FORCE WAS USED DURING DELIVERY. IT WAS REPORTED THAT STENT DISLODGEMENT OCCURRED DURING REMOVAL FOLLOWING A FAILED DELIVERY. THE LESION IN THE MID RCA WAS BEING WORKED ON USING A 7FR AL.75 GUIDE CATHETER AND A NON MEDTRONIC WIRE. AFTER THE LESION WAS PREDILATED USING A NON MEDTRONIC 3.5 X 15 MM BALLOON, SHOCKWAVE WAS USED TO TREAT THE LESION. WHEN ATTEMPTING TO CROSS THE LESION WITH THE ONYX FRONTIER STENT, IT WOULD NOT FULLY CROSS THE LESION. WHEN PULLING THE STENT BACK INTO THE GUIDE CATHETER, IT SHEARED OFF THE BALLOON. THE DISLODGED STENT WAS NOT REMOVED. A 3.5 X 8 MM NON MEDTRONIC BALLOON WAS PLACED INSIDE THE UNDEPLOYED STENT AND WAS USED TO SUCCESSFULLY OPEN THE STENT. THE DISLODGED STENT WAS DEPLOYED. IT WAS BELIEVED THAT THERE WAS POSSIBLY EXCESSIVE FORCE USED DUE TO THE CALCIFICATION AND THE NEED TO USE SHOCKWAVE. THE PATIENT WAS REPORTED TO BE FINE. THE PATIENT IS ALIVE WITH NO FURTHER INJURY. ADDITIONAL INFORMATION: ANNEX D CODES ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20568998 | 2024 | XIENCE SKYPOINTÃ‚Â¿ | Unknown | ['Product Quality Problem', 'Failure to Advance', 'Low Blood Pressure/ Hypotension', 'Dizziness'] | MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT THE CIRCUMFLEX CORONARY ARTERY WITH CALCIFICATION AND 70% STENOSIS. THE 2.5X23 MM XIENCE SKYPOINT STENT DELIVERY SYSTEM (SDS) FAILED TO CROSS THE LESION DUE TO CALCIFICATION AND BECAME DAMAGED DURING THE PROCEDURE. ANOTHER STENT WAS USED TO COMPLETE THE PROCEDURE. DUE TO THE INCREASED LENGTH OF THE PROCEDURE THE PATIENT WAS HOSPITALIZED FOR AN EXTRA DAY DUE TO DIZZINESS AND HYPOTENSION. TREATMENT, IF ANY, WAS NOT REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED.
20570875 | 2024 | XIENCE SKYPOINTÃ‚Â¿ | Female | ['Failure to Advance', 'Material Deformation', 'Vascular Dissection'] | IT WAS REPORTED THE PROCEDURE WAS TO TREAT A LESION IN THE RIGHT CORONARY ARTERY (RCA) WITH MODERATE CALCIFICATION AND MODERATE TORTUOSITY. THE 3.5X38MM XIENCE SKYPOINT STENT DELIVERY SYSTEM (SDS) ATTEMPTED TO CROSS THE TARGET LESION; HOWEVER, FAILED TO CROSS DUE TO THE ANATOMY. THE DISTAL EDGE OF STENT LOOKED DAMAGED AND A DISSECTION WAS NOTED. THEREFORE, A SHOCKWAVE THERAPY AND BALLOONING FOLLOWED BY IMPLANTATION OF A 3X38MM XIENCE SKYPOINT STENT WAS USED TO TREAT THE DISSECTION. THERE WAS NO ADVERSE PATIENT SEQUELA. NO ADDITIONAL INFORMATION WAS PROVIDED. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. IN THIS CASE, IT IS LIKELY THAT THE STENT DELIVERY SYSTEM (SDS) INTERACTED WITH THE MODERATE CALCIFICATION AND MODERATE TORTUOSITY DURING ADVANCEMENT, AS RESISTANCE WITH THE LESION WAS NOTED, RESULTING IN THE REPORTED FAILURE TO ADVANCE, AND SUBSEQUENT MATERIAL DEFORMATION. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED ISSUE AND TREATMENT APPEARS TO BE RELATED TO OPERATIONAL CONTEXT. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THE REPORTED PATIENT EFFECT OF DISSECTION IS LISTED IN THE XIENCE SKYPOINT EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE.
20596932 | 2024 | ONYX FRONTIER | Male | ['Device Dislodged or Dislocated', 'Positioning Problem', 'Device Embedded In Tissue or Plaque', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. DURING A PROCEDURE AN ATTEMPT WAS MADE TO USE ONE ONYX FRONTIER CORONARY DRUG ELUTING STENT TO TREAT A MODERATELY TORTUOUS, MODERATELY CALCIFIED LESION LOCATED IN THE CIRCUMFLEX (CX) ARTERY / OBTUSE MARGINAL (OM). THE DEVICE WAS INSPECTED WITH NO ISSUES NOTED. NEGATIVE PREP WAS PERFORMED WITH NO ISSUES NOTED. THE LESION WAS PRE-DILATED. THE DEVICE DID NOT PASS THROUGH A PREVIOUSLY DEPLOYED STENT. RESISTANCE WAS NOT ENCOUNTERED WHEN ADVANCING THE DEVICE. EXCESSIVE FORCE WAS NOT USED DURING DELIVERY. IT WAS REPORTED THAT THE STENT FAILED TO CROSS THE LESION AND DISLODGED. IT WAS STATED THAT THE STENT WAS ADVANCED TO THE CIRCUMFLEX AND THE 1ST OM WAS THE TARGET LESION BUT THE STENT DID NOT CROSS. THE DEVICE WAS REMOVED. A 6F TELESCOPE CATHETER WAS INSERTED AND THE ORIGINAL STENT CATHETER WAS GOING TO BE RE-INSERTED BUT IT WAS DISCOVERED THAT THE STENT WAS NO LONGER ON THE BALLOON. THE LOCATION OF THE UNDEPLOYED STENT WAS IN THE CIRCUMFLEX ARTERY. A WIRE WAS INSERTED PASSED DISLODGED, UNDEPLOYED STENT IN THE CIRCUMFLEX ARTERY. A 2.0X15 MM BALLOON WAS ADVANCED TO THE UNDEPLOYED STENT AND INFLATED. THE BALLOON CATHETER WAS REMOVED. ANOTHER ONYX FRONTIER STENT (ONYXNG22512UX , LOT NUMBER 0012151821) WAS SUCCESSFULLY DEPLOYED AND CRUSHED THE DISLODGED STENT AGAINST THE ARTERIAL WALL AND COMPLETELY COVEREDTHE DISLODGED STENT. NO FURTHER PATIENT INJURY WAS REPORTED. ADDITIONAL INFORMATION: RESISTANCE WAS NOT NOTED DURING WITHDRAWAL OF THE DEVICE AND EXCESSIVE FORCE WAS NOT USED. THERE WERE NO ISSUES EXPERIENCED WHEN USING THE 6F TELESCOPE GEC. CORRECTION: IME CODE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20600022 | 2024 | XIENCE SKYPOINTÃ‚Â¿ | Unknown | ['Device Dislodged or Dislocated', 'No Clinical Signs, Symptoms or Conditions'] | THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. BASED ON THE INFORMATION PROVIDED, A DEFINITIVE CAUSE FOR THE REPORTED DEVICE DISLODGED OR DISLOCATED WAS UNDETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT WAS LIKELY DUE TO OPERATIONAL CIRCUMSTANCES. IT WAS REPORTED THAT THE STENT DELIVERY SYSTEM (SDS) WAS BEING ADVANCE IN THE LESION WHEN THE STENT DISLODGED FROM THE SDS. FACTORS THAT MAY CONTRIBUTE TO STENT DISLODGEMENT INCLUDE, BUT ARE NOT LIMITED TO, IMPROPER OR INADEQUATE CRIMPING DURING MANUFACTURE, INADVERTENT MISHANDLING DURING PREPARATION, MANIPULATION AGAINST RESISTANCE, INTERACTION WITH CHALLENGING ANATOMY, OR INTERACTION WITH ACCESSORY DEVICES. BASED ON THE INFORMATION PROVIDED AND BECAUSE THE DEVICE WAS NOT RETURNED FOR ANALYSIS, IT IS UNKNOWN WHAT MAY HAVE CAUSED THE REPORTED DEVICE DISLODGED OR DISLOCATED. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT THE LEFT MAIN CORONARY ARTERY WITH MILD CALCIFICATION AND 70% STENOSIS. THE 3.5X23 XIENCE SKYPOINT STENT DELVIERY SYSTEM (SDS) WAS BEING ADVANCED TO THE LESION WHEN THE STENT DISLODGED AND CAME OFF OF THE DELIVERY SYSTEM IN THE LEFT MAIN. A SNARE WAS USED TO RETRIEVE THE STENT. THERE WERE NO ADVERSE PATIENT SEQUELA. ANOTHER XIENCE SDS WAS USED TO COMPLETE THE PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED.
20609129 | 2024 | ONYX FRONTIER | Unknown | ['Device Dislodged or Dislocated', 'Material Deformation', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. AN ATTEMPT WAS MADE TO USE TWO ONYX FRONTIER CORONARY DRUG ELUTING STENTS TO TREAT A LESION. IT WAS REPORTED THAT STENT DEFORMATION OCCURRED IN VIVO DURING POSITIONING/ADVANCEMENT FOR THE FIRST DEVICE (ONYXNG25038UX). THERE WAS A STENT MALFUNCTION OF THE STRUT. IT WAS FURTHER REPORTED THAT STENT DISLODGEMENT ALSO OCCURRED DURING DELIVERY TO/AT LESION. RESISTANCE WAS NOT NOTED DURING WITHDRAWAL OF THE DEVICE, AND EXCESSIVE FORCE WAS NOT USED. THE DISLODGED STENT WAS REMOVED WITH A SNARE. THE SECOND ONYX FRONTIER STENT (ONYXNG25026UX) WAS THEN USED. THIS DEVICE WAS INSPECTED WITH NO ISSUES. IT WAS REPORTED THAT THERE WAS A STENT MALFUNCTION OF THE STRUT. THE STENT WAS REMOVED FROM THE PATIENT. THE PATIENT IS ALIVE WITH NO FURTHER INJURY. ADDITIONAL INFORMATION: LOT NUMBER UPDATED. IT WAS LATER CONFIRMED THAT THE ONYX FRONTIER STENT ONYXNG25026UX DEFORMED AND DISLODGED AND THAT ONYX FRONTIER STENT ONYXNG25038UX DEFORMED ONLY. PRODUCT ANALYSIS: THE DEVICE WAS RETURNED FOR EVALUATION. THE DISLODGED STENT WAS RECEIVED FOR ANALYSIS. THE DELIVERY SYSTEM WAS NOT RECEIVED ALONGSIDE THE DISLODGED STENT. BIOLOGICAL MATERIAL WAS TRAPPED IN THE STENT. STENT WAS STRETCHED AND DEFORMED. NO OTHER DAMAGE EVIDENT TO THE REMAINDER OF THE DEVICE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20610510 | 2024 | SYNERGY XD | Unknown | ['Failure to Advance', 'Device Dislodged or Dislocated', 'Foreign Body In Patient'] | IT WAS REPORTED THAT STENT DISLODGEMENT OCCURRED. THE TARGET LESION WAS LOCATED IN THE LEFT ANTERIOR DESCENDING ARTERY. A 3.00 X 32MM SYNERGY XD DRUG-ELUTING STENT WAS ADVANCED FOR TREATMENT BUT HAD DIFFICULTY CROSSING THE LESION. THE STENT GOT CAUGHT ON THE PREVIOUS STENT IN THE DIAGONAL BRANCH, DISLODGED, AND FLOATED DOWN TO THE ILIAC. THE PHYSICIAN WAS UNABLE TO SNARE THE STENT AND A VASCULAR SURGEON WAS CALLED. THE VASCULAR SURGEON SAID WHERE THE STENT LANDED WOULD BE OKAY SO THE PROCEDURE WAS COMPLETED USING AN ALTERNATE METHOD. THERE WERE NO PATIENT COMPLICATIONS AND THE PATIENT FULLY RECOVERED. IT WAS REPORTED THAT STENT DISLODGEMENT OCCURRED. THE TARGET LESION WAS LOCATED IN THE LEFT ANTERIOR DESCENDING ARTERY. A 3.00 X 32MM SYNERGY XD DRUG-ELUTING STENT WAS ADVANCED FOR TREATMENT BUT HAD DIFFICULTY CROSSING THE LESION. THE STENT GOT CAUGHT ON THE PREVIOUS STENT IN THE DIAGONAL BRANCH, DISLODGED, AND FLOATED DOWN TO THE ILIAC. THE PHYSICIAN WAS UNABLE TO SNARE THE STENT AND A VASCULAR SURGEON WAS CALLED. THE VASCULAR SURGEON SAID WHERE THE STENT LANDED WOULD BE OKAY SO THE PROCEDURE WAS COMPLETED USING AN ALTERNATE METHOD. THERE WERE NO PATIENT COMPLICATIONS AND THE PATIENT FULLY RECOVERED. IT WAS FURTHER REPORTED THAT RESISTANCE WAS ENCOUNTERED WHILE PASSING THROUGH THE PREVIOUSLY IMPLANTED DIAGONAL STENT, BUT THE STENT WAS NOT TOUCHED. NO PRESSURE WAS APPLIED TO THE DELIVERY SYSTEM BALLOON PRIOR TO THE STENT BECOMING DISLODGED.
20611705 | 2024 | PROMUS ELITE | Male | ['Difficult to Remove', 'Detachment of Device or Device Component', 'Positioning Problem', 'Activation Failure', 'Obstruction/Occlusion', 'Foreign Body In Patient', 'Reduced Blood Flow'] | IT WAS REPORTED THAT A DETACHMENT ISSUE OCCURRED. THE TARGET LESION WAS LOCATED IN THE MIDDLE THIRD OF A SEVERELY CALCIFIED VESSEL. THE 2.25 X 24 MM PROMUS ELITE DRUG ELUTING STENT WAS IMPLANTED AT 24 ATM TO TREAT THE TARGET LESION. THERE WAS POOR EXPANSION AND POSITION OF THE STENT DUE TO THE SEVERE CALCIFICATION. THE BALLOON WAS DEFLATED, HOWEVER, WHEN THE BALLOON WAS BEING REMOVED FROM THE STENT, IT BECAME TRAPPED AND DETACHED. THE DEVICE WAS ABANDONED RESULTING IN A RESIDUAL LESION OF 90% STENOSIS WITH AN UNSATISFACTORY RESULT AND TIMI 1 FLOW. IT WAS REPORTED THAT A DETACHMENT ISSUE OCCURRED. THE TARGET LESION WAS LOCATED IN THE MIDDLE THIRD OF A SEVERELY CALCIFIED VESSEL. THE 2.25 X 24 MM PROMUS ELITE DRUG ELUTING STENT WAS IMPLANTED AT 24 ATM TO TREAT THE TARGET LESION. THERE WAS POOR EXPANSION AND POSITION OF THE STENT DUE TO THE SEVERE CALCIFICATION. THE BALLOON WAS DEFLATED, HOWEVER, WHEN THE BALLOON WAS BEING REMOVED FROM THE STENT, IT BECAME TRAPPED AND DETACHED. THE DEVICE WAS ABANDONED RESULTING IN A RESIDUAL LESION OF 90% STENOSIS WITH AN UNSATISFACTORY RESULT AND TIMI 1 FLOW. IT WAS FURTHER STATED THAT THE TARGET LESION WAS 99% STENOSED. THE DEVICE FRAGMENT WAS NOT RETRIEVED.
20614537 | 2024 | XIENCE ALPINE | Unknown | ['Break', 'Stretched', 'No Clinical Signs, Symptoms or Conditions'] | IT WAS REPORTED THAT THE PROCEDURE WAS PERFORMED TO TREAT A LESION IN THE LEFT CIRCUMFLEX CORONARY ARTERY WITH MODERATE CALCIFICATION, MILD TORTUOSITY AND 85% STENOSIS. PRE-DILATION WAS PERFORMED WITH AN UNSPECIFIED 2.50X12MM BALLOON. A 3.0X38MM XIENCE ALPINE STENT WAS IMPLANTED. POST DILATION WAS PERFORMED WITH AN UNSPECIFIED 3.0X15MM NON-COMPLIANT BALLOON WHICH WAS INFLATED TO 16 ATMOSPHERES. AN OPTICAL COHERENCE TOMOGRAPHY (OCT) ASSESSMENT REVEALED THE PROXIMAL STRUT WAS FRACTURED AND THE STENT ELONGATED BY 2-3MM. A 3X23MM XIENCE ALPINE STENT WAS USED TO TREAT THE FRACTURED AND ELONGATED STENT HOWEVER, THE STENT ALSO FRACTURED IN THE SAME AREA, AS CONFIRMED BY OCT IMAGING. THERE WERE NO ADVERSE PATIENT SEQUELA. NO ADDITIONAL INFORMATION WAS PROVIDED. MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. THE ADDITIONAL DEVICE REFERENCED IN B5 IS FILED UNDER A SEPARATE MEDWATCH REPORT NUMBER. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. THE INVESTIGATION WAS UNABLE TO DETERMINE A CONCLUSIVE CAUSE FOR THE REPORTED BREAK AND STRETCHED STENT; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. IT WAS REPORTED THAT THE STENT DELIVERY SYSTEM (SDS) STENT WAS IMPLANTED. AN ASSESSMENT WITH AN OPTICAL COHERENCE TOMOGRAPHY REVEALED THE PROXIMAL STENT FRACTURE AND AN ELONGATED STENT. FACTORS THAT MAY CONTRIBUTE TO STRETCHED AND BROKEN STENT INCLUDE, BUT ARE NOT LIMITED TO, INFLATION ABOVE RATED BURST PRESSURE, INTERACTION WITH CHALLENGING ANATOMY, INTERACTION WITH ACCESSORY DEVICES, OR MANIPULATION AGAINST RESISTANCE. BASED ON THE INFORMATION PROVIDED AND BECAUSE THE DEVICE WAS NOT RETURNED FOR ANALYSIS, IT IS UNKNOWN WHAT MAY HAVE CAUSED THE REPORTED STRETCHED AND BROKEN STENT. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE.
20620902 | 2024 | SYNERGY XD | Unknown | ['Device Dislodged or Dislocated', 'Perforation'] | IT WAS REPORTED THAT STENT DISLODGEMENT AND PERFORATION OCCURRED REQUIRING ADDITIONAL INTERVENTION. THE PATIENT UNDERWENT PERCUTANEOUS CORONARY INTERVENTION. THE TARGET LESION WAS LOCATED IN THE LEFT ANTERIOR DESCENDING ARTERY. AN IVUS WAS USED TO MEASURE THE DISTAL REFERENCE OF 2.5 MM, AND PRE-DILATATION WAS PERFORMED WITH A 3.0 MM SHOCKWAVE BALLOON. AFTER A 3.00 X 48MM SYNERGY XD DRUG-ELUTING STENT WAS PLACED, MULTIPLE POST-DILATATIONS WERE PERFORMED WITH A 4.0 MM, 4.5 MM AND 5.5 MM NON-BOSTON SCIENTIFIC BALLOON CATHETER. HOWEVER, DURING THE FINAL INFLATION, THE STENT SEEMED TO BE FLOATING IN THE ARTERY AND A SMALL PERFORATION WAS NOTED. IN RESPONSE TO THE PERFORATION, A COVERED STENT WAS PLACED, AND BALLOON TAMPONADE WAS APPLIED, WHICH SUCCESSFULLY DIMINISHED THE PERFORATION. THE PATIENT WAS STABLE POST-PROCEDURE.
20621301 | 2024 | PROMUS PREMIER SELECT | Unknown | ['Difficult to Remove', 'Failure to Advance', 'Material Deformation', 'No Clinical Signs, Symptoms or Conditions'] | (B)(6). IT WAS REPORTED THAT REMOVAL DIFFICULTY OCCURRED. THE 90-95% STENOSED TARGET LESION WAS LOCATED IN THE SEVERELY TORTUOUS AND SEVERELY CALCIFIED DISTAL OBTUSE MARGINAL ARTERY. FOLLOWING PRE-DILATION WITH AN NC BALLOON AND AN NC SCORING BALLOON AT 22 ATM, A 38 X 3.00 PROMUS PREMIER SELECT DRUG-ELUTING STENT (DES) WAS INTRODUCED THROUGH A 3.5 6F NON-BOSTON SCIENTIFIC (NON-BSC) GUIDE CATHETER. THE 38 X 3.00 DES COULD NOT PASS THE MID OM. THE DEVICE WAS SNARED INTO THE GUIDE CATHETER FOR REMOVAL AND ONCE OUTSIDE THE PATIENT, STENT DAMAGE WAS NOTICED. A 28 X 3.00 PROMUS PREMIER SELECT DES WAS INTRODUCED BUT FAILED TO ADVANCE PAST THE OSTIAL LESION. THE 28 X 3.00 DES WAS SNARED FOR REMOVAL INTO THE GUIDE CATHETER, BUT ALTHOUGH IT COULD ENTER THE GUIDE CATHETER, IT COULD NOT FULLY EXIT IT. THIS STENT WAS ALSO DAMAGED. THE DECISION WAS MADE TO REPLACE THE GUIDE CATHETER WITH A MACH 1 JL 4.0 7F AND USE A 6F GUIDEZILLA FOR SUPPORT. A 28 X 2.75 PROMUS PREMIER DES WAS INTRODUCED BUT WAS STILL DAMAGED AND COULD NOT CROSS THE MID OM DUE TO TORTUOSITY AND CALCIFICATION IN THE LEFT CIRCUMFLEX ARTERY. THERE WERE NO PATIENT COMPLICATIONS REPORTED, AND THE PATIENT WAS STABLE FOLLOWING THE PROCEDURE. E1 INITIAL REPORTER ADDRESS 1: (B)(6) E1. INITIAL REPORTER PHONE: B5 DESCRIBE EVENT OR PROBLEM: CORRECTED. IT WAS REPORTED THAT REMOVAL DIFFICULTY OCCURRED. THE 90-95% STENOSED TARGET LESION WAS LOCATED IN THE SEVERELY TORTUOUS AND SEVERELY CALCIFIED DISTAL OBTUSE MARGINAL ARTERY. FOLLOWING PRE-DILATION WITH AN NC BALLOON AND AN NC SCORING BALLOON AT 22 ATM, A 38 X 3.00 PROMUS PREMIER SELECT DRUG-ELUTING STENT (DES) WAS INTRODUCED THROUGH A 3.5 6F NON-BOSTON SCIENTIFIC (NON-BSC) GUIDE CATHETER. THE 38 X 3.00 DES COULD NOT PASS THE MID OM. THE DEVICE WAS SNARED INTO THE GUIDE CATHETER FOR REMOVAL AND ONCE OUTSIDE THE PATIENT, STENT DAMAGE WAS NOTICED. A 28 X 3.00 PROMUS PREMIER SELECT DES WAS INTRODUCED BUT FAILED TO ADVANCE PAST THE OSTIAL LESION. THE 28 X 3.00 DES WAS SNARED FOR REMOVAL INTO THE GUIDE CATHETER, BUT ALTHOUGH IT COULD ENTER THE GUIDE CATHETER, IT COULD NOT FULLY EXIT IT. THIS STENT WAS ALSO DAMAGED. THE DECISION WAS MADE TO REPLACE THE GUIDE CATHETER WITH A MACH 1 JL 4.0 7F AND USE A 6F GUIDEZILLA FOR SUPPORT. A 28 X 2.75 PROMUS PREMIER DES WAS INTRODUCED BUT WAS STILL DAMAGED AND COULD NOT CROSS THE MID OM DUE TO TORTUOSITY AND CALCIFICATION IN THE LEFT CIRCUMFLEX ARTERY. THERE WERE NO PATIENT COMPLICATIONS REPORTED, AND THE PATIENT WAS STABLE FOLLOWING THE PROCEDURE. IT WAS ALSO REPORTED THE PROCEDURE WAS COMPLETED WITH A DRUG COATED BALLOON AND WITHOUT USING A STENT.
20621814 | 2024 | PROMUS PREMIER SELECT | Unknown | ['Difficult to Remove', 'Failure to Advance', 'Material Deformation', 'No Clinical Signs, Symptoms or Conditions'] | E1. INITIAL REPORTER ADDRESS 1: (B)(6). E1. INITIAL REPORTER PHONE: (B)(6). IT WAS REPORTED THAT REMOVAL DIFFICULTY OCCURRED. THE 90-95% STENOSED TARGET LESION WAS LOCATED IN THE SEVERELY TORTUOUS AND SEVERELY CALCIFIED DISTAL OBTUSE MARGINAL ARTERY. FOLLOWING PRE-DILATION WITH AN NC BALLOON AND AN NC SCORING BALLOON AT 22 ATM, A 38 X 3.00 PROMUS PREMIER SELECT DRUG-ELUTING STENT (DES) WAS INTRODUCED THROUGH A 3.5 6F NON-BOSTON SCIENTIFIC (NON-BSC) GUIDE CATHETER. THE 38 X 3.00 DES COULD NOT PASS THE MID OM. THE DEVICE WAS SNARED INTO THE GUIDE CATHETER FOR REMOVAL AND ONCE OUTSIDE THE PATIENT, STENT DAMAGE WAS NOTICED. A 28 X 3.00 PROMUS PREMIER SELECT DES WAS INTRODUCED BUT FAILED TO ADVANCE PAST THE OSTIAL LESION. THE 28 X 3.00 DES WAS SNARED FOR REMOVAL INTO THE GUIDE CATHETER, BUT ALTHOUGH IT COULD ENTER THE GUIDE CATHETER, IT COULD NOT FULLY EXIT IT. THIS STENT WAS ALSO DAMAGED. THE DECISION WAS MADE TO REPLACE THE GUIDE CATHETER WITH A MACH 1 JL 4.0 7F AND USE A 6F GUIDEZILLA FOR SUPPORT. A 28 X 2.75 PROMUS PREMIER DES WAS INTRODUCED BUT WAS STILL DAMAGED AND COULD NOT CROSS THE MID OM DUE TO TORTUOSITY AND CALCIFICATION IN THE LEFT CIRCUMFLEX ARTERY. THE PROCEDURE WAS COMPLETED WITH A DRUG COATED BALLOON AND WITHOUT USING A STENT. THERE WERE NO PATIENT COMPLICATIONS REPORTED, AND THE PATIENT WAS STABLE FOLLOWING THE PROCEDURE.
20622170 | 2024 | PROMUS PREMIER SELECT | Unknown | ['Difficult to Remove', 'Failure to Advance', 'Material Deformation', 'No Clinical Signs, Symptoms or Conditions'] | IT WAS REPORTED THAT REMOVAL DIFFICULTY OCCURRED. THE 90-95% STENOSED TARGET LESION WAS LOCATED IN THE SEVERELY TORTUOUS AND SEVERELY CALCIFIED DISTAL OBTUSE MARGINAL ARTERY. FOLLOWING PRE-DILATION WITH AN NC BALLOON AND AN NC SCORING BALLOON AT 22 ATM, A 38 X 3.00 PROMUS PREMIER SELECT DRUG-ELUTING STENT (DES) WAS INTRODUCED THROUGH A 3.5 6F NON-BOSTON SCIENTIFIC (NON-BSC) GUIDE CATHETER. THE 38 X 3.00 DES COULD NOT PASS THE MID OM. THE DEVICE WAS SNARED INTO THE GUIDE CATHETER FOR REMOVAL AND ONCE OUTSIDE THE PATIENT, STENT DAMAGE WAS NOTICED. A 28 X 3.00 PROMUS PREMIER SELECT DES WAS INTRODUCED BUT FAILED TO ADVANCE PAST THE OSTIAL LESION. THE 28 X 3.00 DES WAS SNARED FOR REMOVAL INTO THE GUIDE CATHETER, BUT ALTHOUGH IT COULD ENTER THE GUIDE CATHETER, IT COULD NOT FULLY EXIT IT. THIS STENT WAS ALSO DAMAGED. THE DECISION WAS MADE TO REPLACE THE GUIDE CATHETER WITH A MACH 1 JL 4.0 7F AND USE A 6F GUIDEZILLA FOR SUPPORT. A 28 X 2.75 PROMUS PREMIER DES WAS INTRODUCED BUT WAS STILL DAMAGED AND COULD NOT CROSS THE MID OM DUE TO TORTUOSITY AND CALCIFICATION IN THE LEFT CIRCUMFLEX ARTERY. THE PROCEDURE WAS COMPLETED WITH A DRUG COATED BALLOON AND WITHOUT USING A STENT. THERE WERE NO PATIENT COMPLICATIONS REPORTED, AND THE PATIENT WAS STABLE FOLLOWING THE PROCEDURE. E1 INITIAL REPORTER ADDRESS 1: (B)(6). E1. INITIAL REPORTER PHONE: (B)(6).
20628273 | 2024 | RESOLUTE ONYX RX | Female | ['Device Dislodged or Dislocated', 'Insufficient Information', 'No Clinical Signs, Symptoms or Conditions'] | AN ATTEMPT WAS MADE TO USE ONE RESOLUTE ONYX CORONARY DRUG ELUTING STENT TO TREAT A MODERATELY TORTUOUS, SEVERELY CALCIFIED LESION WITH 90% STENOSIS IN THE PROXIMAL CIRCUMFLEX (CX) ARTERY. THE DEVICE WAS INSPECTED WITH NO ISSUES. NEGATIVE PREP WAS PERFORMED WITH NO ISSUES. THE LESION WAS PRE-DILATED. THE DEVICE DID NOT PASS THROUGH A PREVIOUSLY DEPLOYED STENT. RESISTANCE WAS ENCOUNTERED WHEN ADVANCING THE DEVICE. EXCESSIVE FORCE WAS USED. IT WAS REPORTED THAT STENT DISLODGEMENT OCCURRED DURING INSERTION INTO THE HOMEOSTATIC VALVE. IT WAS STATED THAT THE STENT PASSED THROUGH THE HEMOSTATIC VALVE AND DISLODGED INSIDE THE PATIENT IN THE PROXIMAL VESSEL. THE PATIENT IS ALIVE WITH NO INJURY. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: THE STENT WAS REMOVED BY SNARING TECHNIQUE USING TWO WIRES. THE PATIENT IS ALIVE WITH NO FURTHER INJURY. B. ADVERSE EVENT OR PRODUCT, 2. OUTCOME ATTRIBUTED TO ADVERSE EVENT, 1. TYPE OF REPORTABLE EVENT UPDATED. ANNEX E, F AND D CODES UPDATED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20639965 | 2024 | ONYX TRUCOR | Male | ['Entrapment of Device', 'Deformation Due to Compressive Stress', 'Device Dislodged or Dislocated', 'Material Deformation', 'Insufficient Information', 'No Clinical Signs, Symptoms or Conditions'] | AN ATTEMPT WAS MADE TO USE ONE ONYX TRUCOR CORONARY DRUG-ELUTING STENT TO TREAT A MILDLY CALCIFIED, MILDLY TORTUOUS LESION IN THE MID LEFT ANTERIOR DESCENDING (LAD) ARTERY. THE DEVICE WAS INSPECTED WITH NO ISSUES NOTED. THE LESION WAS PRE-DILATED. THE DEVICE DID NOT PASS THROUGH A PREVIOUSLY DEPLOYED STENT. RESISTANCE WAS NOT ENCOUNTERED WHEN ADVANCING THE DEVICE. EXCESSIVE FORCE WAS NOT USED. IT WAS REPORTED THAT STENT DEFORMATION OCCURRED IN VIVO DURING POSITIONING/ADVANCEMENT. THE PATIENT IS ALIVE WITH NO INJURY. PLEASE NOTE THAT THIS DEVICE (ONYX TRUCOR) IS NOT MARKETED IN THE UNITED STATES; HOWEVER, THE DEVICE IS DEEMED THE SAME AS THE UNITED STATES-MARKETED DEVICE (RESOLUTE ONYX). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. PRODUCT ANALYSIS SUMMARY: THE STENT WAS NOT PRESENT ON THE BALLOON AND DID NOT RETURN FOR ANALYSIS. THE BALLOON FOLDS WERE OPEN ON DEVICE RETURN. CRIMP IMPRES SIONS WERE VISIBLE ON THE EXPOSED BALLOON SURFACE. A KINK WAS EVIDENT TO THE HYPOTUBE. DEFORMATION EVIDENT TO THE DISTAL TIP. NO OTHER DAMAGE EVIDENT TO THE REMAINDER OF THE DEVICE. ADDITIONAL INFORMATION: IT WAS ALSO REPORTED THAT STENT DISLODGEMENT OCCURRED DURING POSITIONING AT THE LESION. THE DISLODGED STENT WAS REMOVED. THE DISLODGED STENT WAS WITHDRAWN TO THE GUIDING CATHETER BUT IT GOT STUCK. AN ATTEMPT WAS THEN MADE TO CROSS DISTALLY FROM THE DISLODGED STENT AND INFLATE A BALLOON, THEN WITHDRAW THE WHOLE SYSTEM. THE PATIENT IS ALIVE WITH NO FURTHER INJURY. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: ANNEX D CODE. CORRECTION: ANNEX A CODE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20642180 | 2024 | PROMUS PREMIER | Male | ['Device Dislodged or Dislocated', 'Material Deformation', 'No Clinical Signs, Symptoms or Conditions'] | (B)(6). IT WAS REPORTED THAT STENT DISLODGEMENT OCCURRED REQUIRING SURGICAL INTERVENTION. THE PATIENT EXPERIENCED AN ACUTE MYOCARDIAL INFARCTION. THE TARGET LESION WAS LOCATED IN THE CORONARY ARTERY. FOLLOWING CORONARY ANGIOGRAPHY AND THROMBECTOMY, A 24 X 4.00MM PROMUS PREMIER DRUG-ELUTING STENT WAS SELECTED FOR TREATMENT. UPON ENTERING THE CORONARY ARTERY OSTIUM, IT WAS NOTED THAT THE STENT BECAME DISTORTED AND DEFORMED. DURING WITHDRAWAL, THE STENT WAS DISLODGED IN THE RADIAL ARTERY. A VASCULAR INCISION WAS MADE FOR REMOVAL OF THE STENT. NO PATIENT COMPLICATIONS WERE REPORTED, AND NO FURTHER DETAILS WERE REPORTED IN RELATION TO THIS EVENT.
20652157 | 2024 | ONYX FRONTIER | Male | ['Device Dislodged or Dislocated', 'Positioning Problem', 'Device Embedded In Tissue or Plaque', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. AN ATTEMPT WAS MADE TO USE ONE ONYX FRONTIER CORONARY DRUG ELUTING STENT (DES) TO TREAT A MODERATELY TORTUOUS, SEVERELY CALCIFIED LESION WITH 50% STENOSIS IN THE PROXIMAL RIGHT CORONARY ARTERY (RCA). THE PATIENT WAS IN ATRIAL FIBRILLATION AT THE START OF THE PROCEDURE. THE DEVICE WAS INSPECTED WITH NO ISSUES. NEGATIVE PREP WAS NOT PERFORMED. THE LESION WAS PRE-DILATED MULTIPLE TIMES WITH A 1.5X12MM NON-MEDTRONIC (MDT) BALLOON AND A 2.25X20MM NON-MDT BALLOON. THE DEVICE DID NOT PASS THROUGH A PREVIOUSLY DEPLOYED STENT. RESISTANCE WAS ENCOUNTERED WHEN ADVANCING THE DEVICE. EXCESSIVE FORCE WAS NOT USED DURING DELIVERY. A LEFT FEMORAL ARTERY APPROACH WAS USED. IT WAS REPORTED THAT STENT DISLODGEMENT OCCURRED DURING DELIVERY TO/AT LESION. AN ATTEMPT WAS MADE TO DELIVER THE STENT DISTALLY, BUT THERE WAS TROUBLE DELIVERING THE STENT. THE STENT WAS BROUGHT BACK TO THE GUIDE AND WAS GOING TO BE TAKEN BACK DOWN, BUT IT WAS FELT THAT THE STENT CAME OFF THE BALLOON A LITTLE. WHEN THE DEVICE WAS EXAMINED, IT WAS POSSIBLE TO TELL THAT THE STENT WAS A LITTLE OFF THE BALLOON DISTALLY. THE DECISION WAS MADE TO INFLATE THE STENT BALLOON AT THE CURRENT LOCATION. THE STENT BALLOON WAS DEPLOYED AT 10 ATM FOR 10 SECONDS. THE STENT BALLOON WAS REMOVED, AND MULTIPLE DIFFERENT SIZED NON-MDT BALLOONS WERE DELIVERED TO THE UNDEPLOYED PORTION OF THE STENT FOR POST-DILATION. THE UNDEPLOYED STENT WAS DILATED WITH THE OTHER BALLOONS. A 2.25X38MM NON-MDT STENT WAS DELIVERED DISTALLY AND DEPLOYED AT 20 ATM FOR 20 SECONDS. ANOTHER 2.25X38MM NON-MDT STENT FAILED TO CROSS THE LESION. FURTHER BALLOON INFLATION WAS PERFORMED AND THEN THE NON-MDT STENT WAS DELIVERED DISTALLY AND DEPLOYED AT 14 ATM AND 20 ATM FOR 20 SECONDS. A 3.25X6MM EUPHORA CORONARY BALLOON WAS USED AND INFLATED MULTIPLE TIMES THROUGHOUT THE CASE, INCLUDED ONE INFLATION TO 24 ATM FOR 20 SECONDS. MULTIPLE DIFFERENT SIZED NON-MDT BALLOONS WERE USED THROUGHOUT THE CASE. THE FINAL RESULT LOOKED GREAT. THE PATIENT IS ALIVE WITH NO FURTHER INJURY. IMAGE ANALYSIS: FIVE STILL IMAGES RECEIVED FOR ANALYSIS. IMAGES APPEAR TO BE CAMERA IMAGES TAKEN OF FLOUROSCOPY IMAGES DISPLAYED ON A MONITOR. TIMESTAMPS ARE NOT VISIBLE ON THE IMAGES. STENOSIS CAN BE OBSERVED IN THE PROXIMAL RIGHT CORONARY ARTERY (RCA), HOWEVER THE RCA IS OBSCURED BY STAPLES IN THE IMAGES PROVIDED. POST-DILATION OF DEPLOYED STENTS IN THE RCA CAN BE OBSERVED, AND DEPLOYED STENTS ARE VISIBLE IN THE PROXIMAL AND DISTAL RCA. WIRING OF THE RCA CAN BE OBSERVED, AND IMPROVED VESSEL PATENCY CAN BE OBSERVED IN THE FINAL IMAGE. IMAGES DO NOT CAPTURE THE REPORTED DISPLACEMENT OR POSITIONING DIFFICULTIES MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: ONCE RESISTANCE WAS FELT UPON THE ATTEMPT TO PASS THE STENT TO THE DESIRED LOCATION, IT WAS STOPPED AND PULLED BACK. UPON PULLING STENT BACK TO GUIDE, THE STENT WAS NOTICED TO HAVE DISLODGED AND DECISION MADE TO DEPLOY THE STENT AT THAT LOCATION. THERE WERE NO ISSUES NOTED WITH THE NC EUPHORA BALLOON. ANNEX D CODES. CORRECTION: IT WAS A 3.25X6MM NC EUPHORA BALLOON THAT WAS USED DURING THE PROCEDURE. IME CODE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20652797 | 2024 | XIENCE SIERRAÃ‚Â¿ | Unknown | ['Material Separation', 'Angina', 'Bradycardia'] | MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. IT WAS REPORTED THAT THE PROCEDURE WAS PERFORMED TO TREAT A MODERATELY CALCIFIED AND MILDLY TORTUOUS LESION IN THE LEFT ANTERIOR DESCENDING ARTERY (LAD). A 2.25X33MM XIENCE SIERRA DRUG ELUTING STENT (DES) WAS DEPLOYED AT THE TARGET LESION; HOWEVER, THE PATIENT WAS NOTED TO HAVE ANGINA AND BRADYCARDIA, AND DURING DEPLOYMENT THE STRUT WAS NOTED TO BE FRACTURED AND SEPARATED. ANOTHER SAME SIZED XIENCE SIERRA WAS DEPLOYED TO COVER THE BROKEN STENT AND THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE PATIENT SEQUELA AND NO CLINICALLY SIGNIFICANT DELAY IN THE PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE PRODUCTION RECORDS AND THE CORRECTIVE AND PREVENTIVE ACTIONS (CAPA) DATABASE WAS PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED NO INDICATION OF A LOT SPECIFIC ISSUE. THE REPORTED PATIENT EFFECT OF ANGINA IS LISTED IN THE XIENCE SIERRA EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. A CONCLUSIVE CAUSE FOR THE REPORTED MATERIAL SEPARATION AND PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. IT WAS REPORTED THAT DURING DEPLOYMENT OF THE STENT, A SEPARATION WAS OBSERVED. FACTORS THAT MAY CONTRIBUTE TO A MATERIAL SEPARATION INCLUDE, BUT ARE NOT LIMITED TO, TENSILE STRENGTH, CORROSION, EXCESSIVE FORCE APPLIED TO DEVICE, INTERACTION WITH ACCESSORY DEVICES, AND/OR INTERACTION WITH CHALLENGING ANATOMY. BASED ON THE INFORMATION PROVIDED AND BECAUSE THE DEVICE WAS NOT RETURNED FOR ANALYSIS, IT IS UNKNOWN WHAT MAY HAVE CAUSED THE REPORTED MATERIAL SEPARATION. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE.
20663531 | 2024 | XIENCE SKYPOINTÃ‚Â¿ | Male | ['Difficult to Remove', 'Failure to Advance', 'Device Dislodged or Dislocated', 'No Clinical Signs, Symptoms or Conditions'] | IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT THE LEFT MAIN, LEFT ANTERIOR DESCENDING, CIRCUMFLEX AND BIFURCATION WITH HEAVY CALCIFICATION. THE 4.0X23MM XIENCE SKYPOINT WAS ATTEMPTED TO ADVANCED; HOWEVER, FAILED TO CROSS DUE TO ANATOMY. UPON REMOVAL THE STENT WAS OBSERVED TO BE FLARED. A 3.5X38MM XIENCE SKYPOINT WAS ADVANCED BUT FAILED TO CROSS DUE TO ANATOMY AND THE STENT STRUT WAS NOTED TO BE FLARED. ANOTHER 3.5X38MM XIENCE SKYPOINT MET RESISTANCE WITH ANATOMY AND FAILED TO CROSS. DURING REMOVAL RESISTANCE WAS FELT WITH ANATOMY AND THE STENT DISLODGED. THE STENT WAS REMOVED VIA SNARE. A 3.0X16MM AND 3.5X33MM XIENCE SKYPOINT WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADVERSE PATIENT SEQUELA AND NO CLINICALLY SIGNIFICANT DELAY IN THE PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED. MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED ISSUE AND TREATMENT APPEARS TO BE RELATED TO OPERATIONAL CONTEXT. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE.
20672130 | 2024 | RESOLUTE ONYX RX | Male | ['Activation, Positioning or Separation Problem', 'Material Deformation', 'Angina', 'Non specific EKG/ECG Changes', 'Obstruction/Occlusion', 'Thrombosis/Thrombus'] | ANDREAS Y. ANDREOU. 'SUB-ACUTE STENT THROMBOSIS IN A BIFURCATION LESION: THE DEVIL IS IN THE DETAILS.' CARDIOVASCULAR DIAGNOSIS AND THERAPY, NO. 4, 2024. DOI: 10.21037/CDT-24-183.Â PMID:Â 39263483. PAGES: 735-739. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. A JOURNAL ARTICLE WAS RECEIVED FOR REVIEW TITLED "SUB-ACUTE STENT THROMBOSIS IN A BIFURCATION LESION: THE DEVIL IS IN THE DETAILS". THISÂ JOURNAL ARTICLEÂ WAS A CASE STUDY. THE PATIENT INITIALLYÂ PRESENTED WITH INFERIOR WALL ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION (MI) AND UNDERWENT PERCUTANEOUS CORONARY INTERVENTION (PCI) TO A CULPRIT RIGHT CORONARY ARTERY (RCA)-POSTERIOR DESCENDING ARTERY (PDA) BIFURCATION LESION. BOTH A 3.0MM AND A 2.5MM RESOLUTE ONYX CORONARY DRUG ELUTING STENT WERE IMPLANTED USING AÂ T-AND-PROTRUSION (TAP) TECHNIQUE.Â POST-PCI DUAL ANTIPLATELET THERAPY (DAPT) WAS COMPOSED OF MAINTENANCE DOSES OF ASPIRIN AND CLOPIDOGREL. EMERGENCY CORONARY ANGIOGRAPHYÂ PERFORMED FOUR DAYS LATER FOR RECURRENT ANGINA AND ST-SEGMENT ELEVATION IN THE INFERIOR LEADS REVEALED IN-STENTÂ OCCLUSION JUST PROXIMAL TO THE DISTAL RCA BIFURCATION.Â RECANALIZATION WAS ACHIEVED AFTER ANGIOPLASTY USING A 1.5 MM BALLOON INFLATED ACROSS BOTH OSTIA OF THE BIFURCATION, FOLLOWED BY MANUAL ASPIRATION OF A LARGE THROMBUS FROM THE OSTIUM OF THE PDA.Â SUBS EQUENT INTRAVASCULAR ULTRASOUND (IVUS) IMAGINGÂ FROM THE RCA REVEALED EXCESSIVE PROTRUSION OF THE PDAÂ STENT INSIDE THE DISTAL RCA, RESULTING IN GENERATION OF AÂ LONG NEOCARINA,Â AND DISTORTION OF THE RCA STENTÂ PROXIMAL TO THE BIFURCATION, WITH STENT STRUTS PROTRUDINGÂ INTO THE LUMEN, BEING OVERLAPPED WITH EACH OTHER, ANDÂ INTERSPERSED WITHIN THE SUBACUTE THROMBUS. THERE WAS MALAPPOSITION OF THE PDA STENT. IT IS BELIEVED THAT AN UNDERSIZED BALLOONÂ WAS USED IN THE INDEX PROCEDURE IN THE RCA FOR THE FINAL KISSING BALLOONÂ INFLATION (KBI), THEREBY EXPLAINING THE GENERATION OF STENTÂ DISTORTION.Â A PRO-THROMBOGENIC BLOOD FLOW DISTURBANCE WITH HIGH SHEAR RATE WAS THEREFORE CREATED BY THE DISTORTED STENT AND REDUNDANT NEOCARINA RESULTING IN STENT THROMBOSIS. PCI FOR STENT THROMBUS WAS PERFORMED WITH A SEQUENTIAL KBI TECHNIQUE USING PROPERLY SIZED BALLOONS POSITIONED LATERALLY TO EACH OTHER WITH A GOOD ANGIOGRAPHIC RESULT. THE BALLOONS WEREÂ INFLATED SUCCESSIVELY AT HIGH PRESSURE (20 ATM).Â SUBSEQUENTLY, KBI WAS PERFORMED USING THE SAME BALLOONS POSITIONED LATERALLY TO EACH OTHER, WHICH WERE INFLATED SIMULTANEOUSLY UP TO 8 ATM AND DEFLATED ALSO SIMULTANEOUSLY TO ENSURE THAT A CENTRAL POSITION OF THE NEOCARINA IS ACHIEVED.Â FINAL IVUS IMAGING REVEALED RE-EXPANSION OF THE RCA STENT AND RECONSTRUCTION OF THE NEOCARINA. THE PATIENT'S DAPT WAS CHANGED. ADDITIONAL INFORMATION: BOTH A 3.0X34MM AND A 2.5X18MM RESOLUTE ONYX CORONARY DRUG ELUTING STENT WERE IMPLANTED USING AÂ T-AND-PROTRUSION (TAP) TECHNIQUE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: ANNEX D CODE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20672134 | 2024 | RESOLUTE ONYX RX | Male | ['Material Deformation', 'Angina', 'Non specific EKG/ECG Changes', 'Obstruction/Occlusion', 'Thrombosis/Thrombus'] | AUTHORS: ANDREAS Y. ANDREOU. SUB-ACUTE STENT THROMBOSIS IN A BIFURCATION LESION: THE DEVIL IS IN THE DETAILS. CARDIOVASCULAR DIAGNOSIS AND THERAPY 4 2024. DOI: 10.21037/CDT-24-183. PMID: 39263483. PAGES: 735-739. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. A JOURNAL ARTICLE WAS RECEIVED FOR REVIEW TITLED "SUB-ACUTE STENT THROMBOSIS IN A BIFURCATION LESION: THE DEVIL IS IN THE DETAILS". THISÂ JOURNAL ARTICLEÂ WAS A CASE STUDY. THE PATIENT INITIALLYÂ PRESENTED WITH INFERIOR WALL ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION (MI) AND UNDERWENT PERCUTANEOUS CORONARY INTERVENTION (PCI) TO A CULPRIT RIGHT CORONARY ARTERY (RCA)-POSTERIOR DESCENDING ARTERY (PDA) BIFURCATION LESION. BOTH A 3.0MM AND A 2.5MM RESOLUTE ONYX CORONARY DRUG ELUTING STENT WERE IMPLANTED USING AÂ T-AND-PROTRUSION (TAP) TECHNIQUE.Â POST-PCI DUAL ANTIPLATELET THERAPY (DAPT) WAS COMPOSED OF MAINTENANCE DOSES OF ASPIRIN AND CLOPIDOGREL. EMERGENCY CORONARY ANGIOGRAPHYÂ PERFORMED FOUR DAYS LATER FOR RECURRENT ANGINA AND ST-SEGMENT ELEVATION IN THE INFERIOR LEADS REVEALED IN-STENTÂ OCCLUSION JUST PROXIMAL TO THE DISTAL RCA BIFURCATION.Â RECANALIZATION WAS ACHIEVED AFTER ANGIOPLASTY USING A 1.5 MM BALLOON INFLATED ACROSS BOTH OSTIA OF THE BIFURCATION, FOLLOWED BY MANUAL ASPIRATION OF A LARGE THROMBUS FROM THE OSTIUM OF THE PDA.Â SUBS EQUENT INTRAVASCULAR ULTRASOUND (IVUS) IMAGINGÂ FROM THE RCA REVEALED EXCESSIVE PROTRUSION OF THE PDAÂ STENT INSIDE THE DISTAL RCA, RESULTING IN GENERATION OF AÂ LONG NEOCARINA,Â AND DISTORTION OF THE RCA STENTÂ PROXIMAL TO THE BIFURCATION, WITH STENT STRUTS PROTRUDINGÂ INTO THE LUMEN, BEING OVERLAPPED WITH EACH OTHER, ANDÂ INTERSPERSED WITHIN THE SUBACUTE THROMBUS. THERE WAS MALAPPOSITION OF THE PDA STENT. IT IS BELIEVED THAT AN UNDERSIZED BALLOONÂ WAS USED IN THE INDEX PROCEDURE IN THE RCA FOR THE FINAL KISSING BALLOONÂ INFLATION (KBI), THEREBY EXPLAINING THE GENERATION OF STENTÂ DISTORTION.Â A PRO-THROMBOGENIC BLOOD FLOW DISTURBANCE WITH HIGH SHEAR RATE WAS THEREFORE CREATED BY THE DISTORTED STENT AND REDUNDANT NEOCARINA RESULTING IN STENT THROMBOSIS. PCI FOR STENT THROMBUS WAS PERFORMED WITH A SEQUENTIAL KBI TECHNIQUE USING PROPERLY SIZED BALLOONS POSITIONED LATERALLY TO EACH OTHER WITH A GOOD ANGIOGRAPHIC RESULT. THE BALLOONS WEREÂ INFLATED SUCCESSIVELY AT HIGH PRESSURE (20 ATM).Â SUBSEQUENTLY, KBI WAS PERFORMED USING THE SAME BALLOONS POSITIONED LATERALLY TO EACH OTHER, WHICH WERE INFLATED SIMULTANEOUSLY UP TO 8 ATM AND DEFLATED ALSO SIMULTANEOUSLY TO ENSURE THAT A CENTRAL POSITION OF THE NEOCARINA IS ACHIEVED.Â FINAL IVUS IMAGING REVEALED RE-EXPANSION OF THE RCA STENT AND RECONSTRUCTION OF THE NEOCARINA. THE PATIENT'S DAPT WAS CHANGED. ADDITIONAL INFORMATION: BOTH A 3.0X34MM AND A 2.5X18MM RESOLUTE ONYX CORONARY DRUG ELUTING STENT WERE IMPLANTED USING AÂ T-AND-PROTRUSION (TAP) TECHNIQUE.Â LOT NUMBER PATIENT'S WEIGHT MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: ANNEX D CODE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20673478 | 2024 | XIENCE SIERRAÃ‚Â¿ | Male | ['Failure to Advance', 'Material Deformation', 'Vascular Dissection'] | THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. THE REPORTED PATIENT EFFECT OF DISSECTION IS LISTED IN THE XIENCE SIERRA EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE (IFU) AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED ISSUES ARE RELATED TO OPERATIONAL CIRCUMSTANCES. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. IT WAS REPORTED THAT THE STENT DELIVERY SYSTEM (SDS) INTERACTED WITH A HEAVILY CALCIFIED AND TORTUOUS LESION DURING ADVANCEMENT AND RESISTANCE WAS NOTED, RESULTING IN THE REPORTED FAILURE TO ADVANCE. IN ADDITION, IT IS LIKELY THAT THE MANIPULATION OF THE SDS, WHEN RESISTANCE WAS MET, CONTRIBUTED TO THE REPORTED MATERIAL DEFORMATION. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING. IT WAS REPORTED THAT THE PROCEDURE WAS PERFORMED TO TREAT A HEAVILY CALCIFIED AND TORTUOUS LESION IN THE CIRCUMFLEX ARTERY (CX). A 2.25X23MM XIENCE ALPINE DRUG-ELUTING STENT (DES) FAILED TO CROSS THE LESION AND WAS NOTED TO HAVE BENT STRUTS WHILE ATTEMPTING TO CROSS THE LESION IN THE LEFT POSTERIOR VENTRICULAR BRANCH. A DISSECTION WAS THEN NOTED THEREFORE ANOTHER STENT WAS IMPLANTED TO COVER THE DISSECTION. THERE WERE NO ADVERSE PATIENT SEQUELA AND NO CLINICALLY SIGNIFICANT DELAY IN THE PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED.
20674702 | 2024 | SYNERGY MEGATRON | Unknown | ['Activation Failure', 'No Clinical Signs, Symptoms or Conditions'] | IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT PRESENTED WITH ATHEROSCLEROSIS AND AN ANGIOPLASTY PROCEDURE WAS PERFORMED. THE TARGET LESION WAS LOCATED IN THE PROXIMAL LEFT ANTERIOR DESCENDING ARTERY. A 4.00 X 24MM SYNERGY MEGATRON DRUG-ELUTING STENT WAS ADVANCED BUT FAILED TO FULLY EXPAND AND THE DEPLOYMENT WAS SUBOPTIMAL. THE PROCEDURE WAS COMPLETED USING AN ALTERNATE METHOD. HOWEVER, THE PATIENT DIED. B2: DATE OF DEATH: USED THE FIRST DAY OF THE AWARE DATE MONTH AS THE DATE OF DEATH AS IT WAS UNKNOWN. B3: DATE OF EVENT: USED THE FIRST DAY OF THE AWARE DATE MONTH AS THE EVENT DATE AS IT WAS UNKNOWN. D6A: IMPLANT DATE: USED THE FIRST DAY OF THE AWARE DATE MONTH AS THE IMPLANT DATE AS IT WAS UNKNOWN. B2: DATE OF DEATH: USED THE FIRST DAY OF THE AWARE DATE MONTH AS THE DATE OF DEATH AS IT WAS UNKNOWN. B3: DATE OF EVENT: USED THE FIRST DAY OF THE AWARE DATE MONTH AS THE EVENT DATE AS IT WAS UNKNOWN. D6A: IMPLANT DATE: USED THE FIRST DAY OF THE AWARE DATE MONTH AS THE IMPLANT DATE AS IT WAS UNKNOWN. IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT PRESENTED WITH ATHEROSCLEROSIS AND AN ANGIOPLASTY PROCEDURE WAS PERFORMED. THE TARGET LESION WAS LOCATED IN THE PROXIMAL LEFT ANTERIOR DESCENDING ARTERY. A 4.00 X 24MM SYNERGY MEGATRON DRUG-ELUTING STENT WAS ADVANCED BUT FAILED TO FULLY EXPAND AND THE DEPLOYMENT WAS SUBOPTIMAL. THE PROCEDURE WAS COMPLETED USING AN ALTERNATE METHOD. HOWEVER, THE PATIENT DIED. IT WAS FURTHER REPORTED THAT THE 4.00 X 24MM SYNERGY MEGATRON DRUG-ELUTING STENT WAS NOT THOUGHT TO HAVE CONTRIBUTED TO THE PATIENT DEATH.
20678415 | 2024 | XIENCE ALPINE | Unknown | ['Device Dislodged or Dislocated', 'Foreign Body In Patient'] | MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT A 90% STENOSED LESION IN THE LEFT ANTERIOR DESCENDING (LAD) ARTERY WITH HEAVY CALCIFICATION AND MODERATE TORTUOSITY. AFTER, LESION PREPARATION WITH A CUTTING BALLOON AND ROTABLATION, A XIENCE ALPINE STENT WAS IMPLANTED AT THE TARGET LESION. THEN, THE 2.25X23MM XIENCE ALPINE STENT DELIVERY SYSTEM (SDS) WAS ADVANCED TO THE TARGET LESION; HOWEVER, THE STENT DISLODGED, AND THE DISLODGED STENT WAS NOT ABLE TO BE FOUND AND REMAINED IN THE PATIENT. A THIRD XIENCE ALPINE STENT WAS THEN ATTEMPTED TO BE ADVANCED; HOWEVER, RESISTANCE WAS NOTED WITH THE ANATOMY AND THE SHAFT OF THE THIRD SDS BECAME BENT. THEREFORE, THE THIRD SDS WAS REMOVED FROM THE PATIENT. THERE WAS NO ADVERSE PATIENT SEQUELA AND NO CLINICALLY SIGNIFICANT DELAY REPORTED IN THE PROCEDURE. A FOURTH XIENCE ALPINE STENT WAS IMPLANTED TO COMPLETE THE PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED. A VISUAL INSPECTION ANALYSIS WAS PERFORMED ON THE RETURNED DEVICE. DEVICE DISLODGED OR DISLOCATED WAS CONFIRMED. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. A CONCLUSIVE CAUSE FOR THE REPORTED STENT DISLODGEMENT AND PATIENT EFFECT(S), CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. IT WAS REPORTED THAT THE STENT DELIVERY SYSTEMÂ¿S (SDS) STENT DISLODGED DURING USE. ANALYSIS OF THE RETURNED DEVICE CONFIRMED THE REPORTED STENT DISLODGEMENT. CRIMP MARKS WERE NOTED BETWEEN THE MARKERS, INDICATING THE STENT WAS ORIGINALLY PLACED CORRECTLY. FACTORS THAT MAY CONTRIBUTE TO A STENT DISLODGEMENT DURING USE INCLUDE, BUT ARE NOT LIMITED TO, FORCED SHEATH REMOVAL, HANDLING OF THE STENT DURING PREPARATION, INTERACTION WITH CHALLENGING ANATOMY, AND/OR INTERACTION WITH ACCESSORY DEVICES. IN THIS CASE, BASED ON THE INFORMATION PROVIDED, AND THE FINDINGS OF RETURN ANALYSIS, IT CANNOT BE DETERMINED WHAT RESULTED IN THE REPORTED STENT DISLODGEMENT. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING.
20684317 | 2024 | SYNERGY XD | Unknown | ['Fracture', 'No Clinical Signs, Symptoms or Conditions'] | IT WAS REPORTED THAT STENT FRACTURE OCCURRED. THE TARGET LESION WAS LOCATED IN THE LEFT ANTERIOR DESCENDING ARTERY. THE 3.00 X 38MM SYNERGY XD DRUG ELUTING STENT WAS IMPLANTED FOR TREATMENT OF THE TARGET LESION. POST IMPLANT, IT WAS NOTED TO BE FRACTURED INTO TWO PARTS DUE TO OVEREXPANSION. ANOTHER STENT WAS IMPLANTED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT COMPLICATIONS, AND THE PATIENT CONDITION POST PROCEDURE WAS STABLE.
20688123 | 2024 | SYNERGY MEGATRON | Unknown | ['Break', 'No Clinical Signs, Symptoms or Conditions'] | B3 - DATE OF EVENT: USED THE FIRST DATE OF THE MONTH OF THE AWARE DATE AS NO EVENT DATE WAS PROVIDED. IT WAS REPORTED THAT SHAFT BREAK OCCURRED REQUIRING INTERVENTION FOR REMOVAL. THE 95% STENOSED TARGET LESION WAS LOCATED IN THE MODERATELY TORTUOUS AND SEVERELY CALCIFIED RIGHT CORONARY ARTERY. A 5.00 X 12MM SYNERGY MEGATRON DRUG-ELUTING STENT WAS ADVANCED TO TREAT THE LESION. THE STENT WAS SUCCESSFULLY DEPLOYED, HOWEVER; UPON REMOVAL, THE PROXIMAL PORTION OF THE SHAFT BROKE AND CAME OUT FROM THE VESSEL. ANOTHER WIRE WAS PLACED DOWN THE VESSEL WITH ANOTHER BALLOON AND THE SHAFT FRAGMENT WAS PINNED AND PULLED OUT OF THE PATIENT COMPLETELY. THE STENT WAS UNAFFECTED, AND THE PROCEDURE WAS COMPLETED WITHOUT ANY PATIENT COMPLICATIONS REPORTED. THE PATIENT STATUS WAS FINE POST PROCEDURE. B3 - DATE OF EVENT: USED THE FIRST DATE OF THE MONTH OF THE AWARE DATE AS NO EVENT DATE WAS PROVIDED. DEVICE EVALUATED BY MFR.: THE DEVICE SYNERGY MEGATRON MR OUS 5.00 X 12MM STENT DELIVERY SYSTEM WAS RETURNED FOR ANALYSIS. VISUAL, TACTILE AND MICROSCOPIC ANALYSIS WAS PERFORMED ON THE DEVICE. VISUAL EXAMINATION IDENTIFIED THAT THE STENT HAD BEEN DETACHED FROM THE DELIVERY SYSTEM AND NOT RETURNED. VISUAL AND TACTILE EXAMINATION IDENTIFIED A HYPOTUBE KINK 54CM DISTAL TO THE DISTAL END OF THE STRAIN RELIEF. A VISUAL AND TACTILE EXAMINATION OF THE OUTER AND INNER LUMEN AND MID-SHAFT SECTION IDENTIFIED A BREAK AT 119.5CM DISTAL TO THE DISTAL END OF THE STRAIN RELIEF. IN ADDITION, MULTIPLE MID SHAFT KINKS WERE NOTED. MICROSCOPIC EXAMINATION IDENTIFIED MULTIPLE HYPOTUBE KINKS ALONG THE LENGTH OF THE HYPOTUBE SHAFT. BUMPER TIP SHOWED NO SIGNS OF DISTAL TIP DAMAGE. THE BALLOON WAS RETURNED WITH BLOOD PRESENT, HOWEVER DUE TO THE BREAK THE DEVICE COULD NOT BE INFLATED. THE DEVICE WAS RETURNED WITHOUT THE STENT ATTACHED. THE STENT OUTER DIAMETER AT THE TIME OF MANUFACTURING WAS WITHIN MAXIMUM CRIMPED STENT PROFILE MEASUREMENT. NO OTHER DEVICE ISSUES WERE IDENTIFIED DURING RETURNED PRODUCT ANALYSIS. IT WAS REPORTED THAT SHAFT BREAK OCCURRED REQUIRING INTERVENTION FOR REMOVAL. THE 95% STENOSED TARGET LESION WAS LOCATED IN THE MODERATELY TORTUOUS AND SEVERELY CALCIFIED RIGHT CORONARY ARTERY. A 5.00 X 12MM SYNERGY MEGATRON DRUG-ELUTING STENT WAS ADVANCED TO TREAT THE LESION. THE STENT WAS SUCCESSFULLY DEPLOYED, HOWEVER; UPON REMOVAL, THE PROXIMAL PORTION OF THE SHAFT BROKE AND CAME OUT FROM THE VESSEL. ANOTHER WIRE WAS PLACED DOWN THE VESSEL WITH ANOTHER BALLOON AND THE SHAFT FRAGMENT WAS PINNED AND PULLED OUT OF THE PATIENT COMPLETELY. THE STENT WAS UNAFFECTED, AND THE PROCEDURE WAS COMPLETED WITHOUT ANY PATIENT COMPLICATIONS REPORTED. THE PATIENT STATUS WAS FINE POST PROCEDURE.
20693001 | 2024 | ONYX FRONTIER | Male | ['Device Dislodged or Dislocated', 'Positioning Problem', 'Device Embedded In Tissue or Plaque', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. AN ATTEMPT WAS MADE TO USE ONE ONYX FRONTIER CORONARY DRUG ELUTING STENT TO TREAT A MODERATELY TORTUOUS, SEVERELY CALCIFIED LESION WITH 99% STENOSIS IN THE PROXIMAL RIGHT CORONARY ARTERY (RCA). THE PATIENT WAS BEING TREATED FOR AN ST-ELEVATION MYOCARDIAL INFARCTION (STEMI). THE DEVICE WAS INSPECTED WITH NO ISSUES. NEGATIVE PREP WAS PERFORMED WITH NO ISSUES. THE LESION WAS PRE-DILATED WITH A 2.5X20MM EUPHORA BALLOON, AND A 2.75MM NON-MEDTRONIC NC BALLOON. THE DEVICE DID NOT PASS THROUGH A PREVIOUSLY DEPLOYED STENT. RESISTANCE WAS ENCOUNTERED WHEN ADVANCING THE DEVICE. IT WAS REPORTED THAT STENT DISLODGEMENT OCCURRED DURING REMOVAL FOLLOWING A FAILED DELIVERY. THE STENT COULD NOT CROSS THE LESION OF THE TOTAL OCCLUSION. UPON REMOVAL, THE STENT CAME OFF THE BALLOON. THE DISLODGED STENT WAS NOT REMOVED. IT WAS POSSIBLE TO OPEN THE DISLODGED STENT IN THE OSTIUM OF THE RCA USING A 2.75X12MM EUPHORA BALLOON. A MEDTRONIC 6F LAUNCHER GUIDE CATHETER AND A MEDTRONIC TELESCOPE GUIDE EXTENSION CATHETER WERE ALSO USED DURING THE PROCEDURE. THE PATIENT IS ALIVE WITH NO FURTHER INJURY. ADDITIONAL INFORMATION: IT WAS POSSIBLE TO OPEN THE DISLODGED STENT IN THE OSTIUM OF THE RCA USING A 2.75X12MM EUPHORA BALLOON WITH NO ISSUES NOTED. A MEDTRONIC 6F LAUNCHER GUIDE CATHETER AND A MEDTRONIC TELESCOPE GUIDE EXTENSION CATHETER WERE ALSO USED DURING THE PROCEDURE WITH NO ISSUES NOTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: ANNEX D CODES. CORRECTION: THERE IS NO COMPLAINT AGAINST ANY OF THE OTHER MEDTRONIC DEVICES USED DURING THE PROCEDURE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20693160 | 2024 | ONYX FRONTIER | Unknown | ['Device Dislodged or Dislocated', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. AN ATTEMPT WAS MADE TO USE ONE ONYX FRONTIER CORONARY DRUG ELUTING STENT TO TREAT A LESION. THE DEVICE WAS INSPECTED WITH NO ISSUES. IT WAS REPORTED THAT STENT DISLODGEMENT OCCURRED DURING DELIVERY TO/AT THE LESION. THE DISLODGED STENT WAS REMOVED WITH AN ENDOVASCULAR SNARE WITHIN THE GUIDE CATHETER. THE PATIENT IS ALIVE WITH NO FURTHER INJURY. ADDITIONAL INFORMATION: AN ATTEMPT WAS MADE TO USE ONE ONYX FRONTIER CORONARY DRUG ELUTING STENT TO TREAT A LESION IN THE LEFT MAIN/OSTIAL CIRCUMFLEX ARTERY. THE DEVICE WAS PREPPED PER IFU WITH NO ISSUES NOTED. THE DEVICE DID NOT PASS THROUGH A PREVIOUSLY DEPLOYED SENT OR CELL OF A STENT. THE STENT DISLODGED INSIDE THE NON-MEDTRONIC GUIDE CATHETER. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: AN ATTEMPT WAS MADE TO USE ONE ONYX FRONTIER CORONARY DRUG ELUTING STENT TO TREAT A LESION IN THE OSTIAL CIRCUMFLEX ARTERY. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20693321 | 2024 | RESOLUTE ONYX RX | Male | ['Inadequacy of Device Shape and/or Size', 'Improper or Incorrect Procedure or Method', 'Hematoma', 'Obstruction/Occlusion', 'Vascular Dissection', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. THE PATIENT WAS ADMITTED TO HOSPITAL DUE TO ACUTE MYOCARDIAL INFARCTION. EMERGENCY CORONARY ANGIOGRAPHY WAS PERFORMED WHICH SHOWED COMPLETE OCCLUSION OF THE PROXIMAL CIRCUMFLEX (CX) BRANCH. ONE RESOLUTE ONYX CORONARY DRUG ELUTING STENT WAS INSERTED AT THE LESION OF THE CX BRANCH AND WAS FULLY EXPANDED. THERE WAS NO CHEST TIGHTNESS OR PAIN AFTER THE OPERATION. ONE DAY POST STENT IMPLANTATION THE ELECTROCARDIOGRAM (ECG) WAS RECHECKED, AND THE ST SEGMENT RETURNED TO THE BASELINE. APPROXIMATELY 11 DAYS POST STENT IMPLANTATION A RE-DO CORONARY ANGIOGRAPHY WAS PERFORMED WHICH INDICATED THAT THE RESOLUTE ONYX STENT IMPLANTED IN THE CX BRANCH WAS OCCLUDED. THE PROXIMAL END OF THE STENT WAS CLOSED, AND IT WAS DIFFICULT TO PASS THE GUIDEWIRE THROUGH THE LESION. THE GUIDEWIRE WENT INTO A FALSE LUMEN. THERE WAS NO PROBLEM WITH THE STENT ITSELF OR THE STENT DELIVERY SYSTEM. THE GUIDEWIRE WAS PASSED THROUGH THE TRUE LUMEN OF THE BLOOD VESSEL, AND ANOTHER STENT WAS ADDED. THE SECOND SURGERY WAS SUCCESSFUL. IT IS BELIEVED THAT THE RESOLUTE ONYX STENT WAS TOO SMALL AND DID NOT FULLY COVER THE PLAQUE, RESULTING IN DELAYED HEMATOMA OF THE VASCULAR DISSECTION PROXIMAL TO THE STENT. THE P PATIENT RECOVERED AND WAS DISCHARGED FROM THE HOSPITAL. PRODUCT ANALYSIS: ONE MP4 VIDEO WAS PROVIDED FOR REVIEW. IMAGE SHOWS INJECTION OF CONTRAST INTO THE LEFT CORONARY SYSTEM. THERE APPEARS TO BE AN OCCLUSION IN THE SIDE BRANCH VESSEL. THERE ALSO APPEARS TO BE A LEVEL OF VESSEL DISSECTION WHERE THE CONTRAST HAS POOLED IN THE DISSECTION IN THE VESSEL LUMEN. NO IMAGES WERE PROVIDED SHOWING THE STENT DEPLOYMENTS OR ANY ATTEMPT TO TREAT THE LESION/OCCLUSION. ROOT CAUSE OF REPORTED ISSUES CANNOT BE DETERMINED FROM THE VIDEO PROVIDED. BUT OCCLUSION AND DISSECTION CAN BE CONFIRMED. ADDITIONAL INFORMATION: THE RESOLUTE ONYX STENT WAS INITIALLY USED TO TREAT THE STENOSIS. THE PATIENT WAS TRANSFERRED TO ANOTHER HOSPITAL. THE GUIDEWIRE USED WAS A NON-MEDTRONIC DEVICE. THE BLOOD VESSELS WERE OPENED AFTER THE SECOND STENT WAS IMPLANTED, AND THE SURGERY WAS COMPLETED. THE RESOLUTE ONYX STENT DID NOT CAUSE THE DISSECTION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: ANNEX D CODE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20694759 | 2024 | XIENCE SIERRAÃ‚Â¿ | Unknown | ['Break', 'No Clinical Signs, Symptoms or Conditions'] | THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE PRODUCTION RECORDS AND THE CORRECTIVE AND PREVENTIVE ACTIONS (CAPA) DATABASE WAS PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED NO INDICATION OF A LOT SPECIFIC ISSUE. THE INVESTIGATION WAS UNABLE TO DETERMINE A CONCLUSIVE CAUSE FOR THE REPORTED BREAK. ADDITIONALLY, THE REPORTED TREATMENT APPEARS TO BE RELATED TO OPERATIONAL CIRCUMSTANCES. IT WAS REPORTED THAT POST IMPLANTATION THE STENT WAS NOTED TO BE BROKEN. FACTORS THAT MAY CONTRIBUTE TO A BREAK INCLUDE, BUT ARE NOT LIMITED TO, TENSILE STRENGTH, CORROSION, OVER INFLATION OF THE STENT, INTERACTION WITH ACCESSORY DEVICES, AND/OR INTERACTION WITH CHALLENGING ANATOMY. IN THIS CASE, IT IS POSSIBLE THAT THE STENT SUSTAINED DAMAGE DURING ADVANCEMENT, RESULTING IN THE BREAK. IT SHOULD BE NOTED THAT NO DAMAGE OR ANOMALIES WERE NOTED TO THE STENT DURING THE EXAMINATION PRIOR TO USE. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. IT WAS REPORTED THAT THE PROCEDURE WAS PERFORMED TO TREAT A LESION IN THE LEFT ANTERIOR DESCENDING CORONARY ARTERY WITH 100% STENOSIS. THE 2.5X33MM XIENCE SIERRA STENT DELIVERY SYSTEM WAS ADVANCED AND THE STENT WAS IMPLANTED; HOWEVER, THE STRUTS WERE NOTED TO BE FRACTURED, YET REMAINED IN ONE PIECE. A 2.5X38MM XIENCE SIERRA STENT WAS USED TO TREAT THE FRACTURED STENT. THERE WERE NO ADVERSE PATIENT SEQUELA. NO ADDITIONAL INFORMATION WAS PROVIDED.
20704933 | 2024 | RESOLUTE ONYX RX | Male | ['Migration or Expulsion of Device', 'Pain', 'Obstruction/Occlusion', 'Thrombosis/Thrombus', 'Insufficient Information'] | JOURNAL ARTICLE: CORONARY STENT MIGRATION TO AORTA WITH SUCCESSFUL RECOVERY DURING AORTOBIFEMORAL BYPASS JOURNAL: JOURNAL OF VASCULAR SURGERY CASES, INNOVATIONS AND TECHNIQUES YEAR: 2024 REF: HTTPS://DOI.ORG/10.1016/J.JVSCIT.2024.101456 B3: DATE OF PUBLICATION MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. A CASE STUDY WAS SUBMITTED FOR REVIEW TITLED "CORONARY STENT MIGRATION TO AORTA WITH SUCCESSFUL RECOVERY DURING AORTOBIFEMORAL BYPASS".Â A PATIENT WITHÂ SEVERE PERIPHERAL ARTERIAL DISEASE AND A KNOWN CHRONIC INFRARENAL AORTIC OCCLUSION CAUSING SEVERE SHORT-DISTANCE CLAUDICATION PRESENTED TO THE CLINIC WITH PRIOR COMPUTED TOMOGRAPHY ANGIOGRAPHY CONFIRMING DISTAL AORTIC OCCLUSION WITHÂ RECONSTITUTION IN THE BILATERAL COMMON FEMORAL ARTERIES.Â HOWEVER, A ROUTINE PREOPERATIVE WORKUP SHOWED ISCHEMIC CARDIOMYOPATHY AND MULTIVESSEL CORONARY DISEASE REQUIRING PERCUTANEOUS CORONARY INTERVENTION (PCI) AND CORONARY STENTING. THE LEFT MAIN CORONARY ARTERY HAD OSTIAL 80% STENOSIS, THE PROXIMAL LEFT ANTERIOR DESCENDING ARTERY (LAD) HAD DIFFUSE 80% TO 90% STENOSIS, AND THE RIGHT CORONARY ARTERY HAD SEVERE 80% OSTIAL STENOSIS. FOR TREATMENT, SIX STENTS WERE ORIGINALLY DEPLOYED. ONEÂ 3.5 X 8MM RESOLUTE ONYX DRUG ELUTING STENT WAS IMPLANTEDÂ IN THE DISTAL OSTIUM OF THE LEFT MAIN CORONARY ARTERY. THREEÂ RESOLUTE ONYX DRUG ELUTING STENTSÂ OF SIZESÂ 3.5 X 22MM, 2.5 X 26MM, AND 2.25 X 18MM WERE IMPLANTED OVERLAPPINGÂ IN THE LAD. TWOÂ RESOLUTE ONYX DRUG ELUTING STENTS OF SIZESÂ 3.0 X 38MM AND 2.75 X30MM. WERE IMPLANTED OVERLAPPINGÂ IN THE RIGHT CORONARY ARTERY.Â HOWEVER, SOME PROXIMAL DISEASE REMAINED IN THE LEFT MAIN CORONARY ARTERY, AND A SECOND MORE PROXIMALÂ 4.0 X 8MMÂ RESOLUTE ONYX DRUG ELUTING STENT WAS DEPLOYED, OVERLAPPING THE ORIGINAL STENT. THE FINAL CORONARY ANGIOGRAM SHOWED EXCELLENT RESULTS, WITH 0% RESIDUAL STENOSIS AND NO EVIDENCE OF DISSECTION OR PERFORATION.Â THE PATIENT COMPLETED THE RECOMMENDED RECOVERY PERIOD OF 3 MONTHS BY CARDIOLOGY AND RETURNED TO THE VASCULAR SURGERY CLINIC FOR OPERATIVE CONSULTATION. THE DECISION WAS MADE TO ATTEMPT TO MANAGE THE PATIENTÂ¿S CLAUDICATION SYMPTOMS WITHOUT SURGERY; HOWEVER, OVER TIME, THE SYMPTOMS BECAME PROGRESSIVELY WORSE, AND, AFTER SEVERAL MONTHS, THE PATIENT REQUESTED OPERATIVE INTERVENTION. A NEW COMPUTED TOMOGRAPHY ANGIOGRAM WAS OBTAINED, DEMONSTRATING A NEW RADIOPAQUE FOREIGN BODY IN THE IMMEDIATE INFRARENAL AORTA AT THE LEVEL OF THE AORTIC OCCLUSION. THIS FOREIGN BODY HAD THE APPEARANCE OF A SMALL-DIAMETER INTRAVASCULAR STENT, SUSPICIOUS FOR A CORONARY STENT THAT HAD EMBOLIZED. OPEN AORTOBIFEMORAL BYPASS WAS, THEREFORE, PERFORMED 11 MONTHS AFTER THE CARDIAC CATHETERIZATION AND PCI PROCEDURE. CLOPIDOGREL WAS HELD FOR 7 DAYS PREOPERATIVELY, WITH 81 MG OF ASPIRIN CONTINUED PERIOPERATIVELY. THE BYPASS SURGERY USED A BIFURCATED DACRON GRAFT, WITH THE PROXIMAL ANASTOMOSIS ORIENTED IN A STANDARD END-TO-END CONFIGURATION.Â DURING EVACUATION OF THE THROMBUS, THE FOREIGN BODY WAS ENCOUNTERED AND CONFIRMED TO BE A CORONARY STENT. THE PATIENT TOLERATEDÂ THE PROCEDURE WELL, AND CARDIOLOGY WAS CONSULTED ON POSTOPERATIVE DAY 0 WITH THIS FINDING. BECAUSE THE PATIENT WAS STABLE, WITH NO SIGNS OF CARDIAC ISCHEMIA, THE DECISION WAS MADE TO PERFORM A DELAYED REPEAT PCI AFTER RECOVERY FROM THE INITIAL INSULT OF SURGERY. FIVE DAYS POSTOPERATIVELY, A REPEAT CORONARY ANGIOGRAM WAS PERFORMED, DEMONSTRATING A PATENT LEFT MAINÂ CORONARY ARTERY STENT DISTALLY WITH AN ABSENT LEFT MAIN CORONARY STENT PROXIMALLY. BECAUSE NO RECURRENT OSTIAL STENOSIS WAS PRESENT, REPEAT STENTING WAS DEEMED UNNECESSARY. NO FURTHER PROCEDURES WERE DONE, AND THE PATIENT REMAINS WITHOUT CHEST PAIN, SHORTNESS OF BREATH, OR WEIGHT GAIN AT 6 MONTHS OF FOLLOW-UP. ADDITIONAL INFORMATION: ANNEX D CODE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20709383 | 2024 | SYNERGY XD | Unknown | ['Failure to Advance', 'Device Dislodged or Dislocated', 'Material Deformation', 'Foreign Body In Patient'] | IT WAS REPORTED THAT DISLODGEMENT OCCURRED REQUIRING ADDITIONAL INTERVENTION. THE TARGET LESION, WHICH WAS 90% STENOSED, WAS LOCATED IN THE MODERATELY TORTUOUS AND SEVERELY CALCIFIED LEFT CIRCUMFLEX ARTERY. A 3.50 X 16MM SYNERGY XD DRUG-ELUTING STENT WAS SELECTED FOR TREATMENT. A NON-BOSTON SCIENTIFIC 6F GUIDE CATHETER AND A NON-BOSTON SCIENTIFIC 0.014-INCH GUIDEWIRE WERE ALSO USED. TWO ATTEMPTS WERE MADE TO DELIVER THE STENT INTO THE CIRCUMFLEX ARTERY; HOWEVER, THE PRESENCE OF MODERATE TORTUOSITY, SEVERE CALCIFICATION, AND A PRIOR STENT IN THE LEFT MAIN ARTERY CAUSED THESE ATTEMPTS TO FAIL. UPON REMOVAL OF THE STENT INFLATION DEVICE, FLUOROSCOPY REVEALED THAT THE STENT HAD INADVERTENTLY MOVED INTO THE LEFT MAIN ARTERY AFTER DETACHING FROM THE BALLOON BEFORE REACHING THE TARGET LESION. THIS CAUSED COMPLICATIONS, AS THE STENT WAS NO LONGER IN THE INTENDED POSITION. THE PHYSICIAN THEN ATTEMPTED TO PASS A BALLOON THROUGH THE UNDEPLOYED STENT, AIMING TO SLIGHTLY INFLATE THE BALLOON IN ORDER TO ENGAGE THE STENT AND PULL IT BACK INTO THE GUIDE CATHETER. HOWEVER, THIS ATTEMPT WAS UNSUCCESSFUL. FLUOROSCOPY SHOWED THAT THE STENT HAD SHORTENED AND WAS NOW POSITIONED IN THE DISTAL LEFT MAIN AND OSTIAL LAD. DUE TO THE DEFORMATION OF THE STENT AND THE INABILITY TO RETRIEVE IT, ADDITIONAL BALLOONS WERE USED TO CRUSH THE STENT AGAINST THE VESSEL WALL, ENSURING CONTINUED BLOOD FLOW. THE PROCEDURE WAS ABORTED, AND THE PATIENT FULLY RECOVERED.
20727305 | 2024 | XIENCE SKYPOINTÃ‚Â¿ | Male | ['Material Deformation', 'Obstruction/Occlusion', 'Thrombosis/Thrombus'] | MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. THE ADDITIONAL DEVICES MENTIONED IN B5 ARE CAPTURED UNDER SEPARATE MEDWATCH NUMBERS. IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT THE LEFT CIRCUMFLEX (LCX) AND LEFT ANTERIOR DESCENDING (LAD) CORONARY ARTERIES. PRE-DILATATION WAS PERFORMED WITH 2.0X15 MM AND 2.5X20 MM NC BALLOONS. A NON-ABBOTT STENT DELIVERY SYSTEM (SDS) COULD NOT BE ADVANCED TO THE LCX SO A 6FR GUIDE CATHETER EXTENSION WAS USED AND THE 2.5X28 MM XIENCE SKYPOINT SDS WAS ADVANCED TO THE LESION AND BEFORE DEPLOYMENT, THE GUIDE CATHETER EXTENSION WAS REMOVED AND A 2.5X15 MM NC BALLOON WAS PLACED IN THE LAD. THEN THE 2.5X28 MM XIENCE SKYPOINT STENT WAS DEPLOYED IN THE LCX AND WAS CRUSHED WITH THE LAD BALLOON. THE 2.25X28 MM XIENCE SKYPOINT WAS DEPLOYED IN THE PROXIMAL LAD AND A 3.0X28 MM XIENCE SKYPOINT DEPLOYED IN THE OSTIAL LEFT MAIN TO LAD. PRE-DILATATION WAS THEN PERFORMED IN THE LCX WITH A 2.0X8 MM COMPLIANT BALLOON AND A 2.25X12 MM NC BALLOON. KISSING BALLOON TECHNIQUE WAS ATTEMPTED BUT THE LCX BALLOON WAS DIFFICULT TO MOVE. THIS BALLOON WAS REMOVED AND A NEW 2.25 NC BALLOON WAS USED TO PERFORM KISSING BALLOON TECHNIQUE WITH THE LAD. DURING THIS TIME IT WAS NOTED POSSIBLE LONGITUDINAL STENT DEFORMATION OF THE STENT IN THE LEFT MAIN AND BALLOON ANGIOPLASTY WAS PERFORMED WITH A 3.75X8 MM NC BALLOON. INTRAVASCULAR ULTRASOUND (IVUS) SHOWED THAT THE OSTIAL AND PROXIMAL LEFT MAIN WERE NOT COVERED DUE TO THE STENT DEFORMATION AND THERE WAS ALSO A POSSIBLE FILLING DEFECT IN THE LEFT MAIN. AGGRASTAT WAS GIVEN AND A 3.5X12 MM XIENCE SKYPOINT STENT WAS PLACED IN THE OSTIAL LEFT MAIN. POST DILATATION WAS PERFORMED WITH THE STENT BALLOON AS WELL AS A 4.0X12 MM NC BALLOON. FINAL ANGIOGRAM SHOWED GOOD FLOW IN THE LEFT MAIN, LAD AND LCX WITH SLIGHT UNDEREXPANSION IN THE PROXIMAL LAD AND PROXIMAL LCX. THERE WERE NO REPORTED ADVERSE PATIENT SEQUELA. NO ADDITIONAL INFORMATION WAS PROVIDED. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. A CONCLUSIVE CAUSE FOR THE REPORTED DIFFICULTIES AND PATIENT EFFECTS, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. IT WAS REPORTED THAT THE STENT DELIVERY SYSTEM (SDS) WAS IMPLANTED HOWEVER STENT DEFORMATION WAS NOTED. FACTORS THAT MAY CONTRIBUTE TO MATERIAL DEFORMATION INCLUDE, BUT ARE NOT LIMITED TO, STENT RECOIL, IMPLANTATION TECHNIQUE, DAMAGE TO THE STENT, OR VESSEL CHARACTERISTICS. BASED ON THE INFORMATION PROVIDED AND BECAUSE THE DEVICE WAS NOT RETURNED FOR ANALYSIS, IT IS UNKNOWN WHAT MAY HAVE CAUSED THE REPORTED MATERIAL DEFORMATION. THE REPORTED PATIENT EFFECT OF THROMBOSIS IS LISTED IN THE XIENCE SKYPOINT EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. A CINE WAS RECEIVED AND REVIEWED. THE REVIEWER CONCLUDED THAT A SERIES OF ANGIOGRAPHIC STILLS OF NON-CONTRAST FILLED LEFT MAIN (LM), LEFT CIRCUMFLEX (LCX), AND LEFT ANTERIOR DESCENDING (LAD) WHERE IT SHOWS STENTS WITHIN EACH ANATOMY UNDERGOING DIFFERENT BALLOON DILATATIONS AND TECHNIQUES OF THE ANATOMY MENTIONED. IT IS UNCLEAR OF WHERE EACH IMAGE REPRESENTS WHERE WE ARE SEQUENTIALLY IN THE CASE. TWO IMAGES REPRESENT A SIMULTANEOUS BALLOON DILATION IN THE LAD AND LCX, NOT RELATED TO THE KISSING BALLOON TECHNIQUE OR THE PROXIMAL OPTIMIZATION TECHNIQUE MENTIONED IN THE REPORT. ANOTHER IMAGE LACKS ANATOMICAL IDENTIFIERS, BUT WHAT APPEARS TO BE POTENTIALLY THE LM WITH BALLOON (IDENTIFIABLE BY BALLOON MARKERS) BETWEEN THE PROXIMAL END OF POTENTIALLY THE LM AND THE PRESUMED GUIDE CATHETER. HERE THE PROXIMAL END OF THE STENT APPEARS MALFORMED BUT CANNOT CONFIRM DUE TO QUALITY OF IMAGE AND THE INABILITY TO VIEW THE ENTIRE STENT. WHILE I AM UNDER THE IMPRESSION THAT THIS IMAGE DOES SHOW A MALFORMATION IN A STENT WITHIN THE ANATOMY, A PROBABLE CAUSE CANNOT BE DETERMINED WITH THE INFORMATION AND MEDIA PROVIDED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. H6 HEALTH EFFECT - CLINICAL CODE: CORRECTION 2422 REMOVED CODE 4440 ADDED. SUBSEQUENT TO THE INITIALLY FILED REPORT, IT WAS REPORTED THAT THE 2.5X28 MM XIENCE SKYPOINT STENT WAS CRUSHED IN A DOUBLE KISSING TECHNIQUE PER STANDARD PROCEDURE. THE SMALL FILLING DEFECT SEEN IN THE 3.0X28 MM XIENCE SKYPOINT STENT WAS POSSIBLE SMALL THROMBUS AND THIS IS THE REASON AGGRASTAT WAS GIVEN. NO ADDITIONAL INFORMATION WAS PROVIDED.
20739262 | 2024 | SYNERGY XD | Unknown | ['Break', 'Foreign Body In Patient', 'Vascular Dissection'] | IT WAS REPORTED THAT SHAFT BREAK OCCURRED AND RESULTED IN AN UNRETRIEVED DEVICE FRAGMENT. A 4.00 X 24 MM SYNERGY XD DRUG-ELUTING STENT WAS SELECTED FOR TREATMENT; HOWEVER, DURING THE PROCEDURE, A SHAFT BREAK OCCURRED. THE DEVICE WAS PARTIALLY EXPLANTED; HOWEVER, A PORTION REMAINS IMPLANTED. A THE PROCEDURE WAS COMPLETED USING AN ALTERNATE DEVICE, AND NO PATIENT COMPLICATIONS WERE REPORTED. B2. OUTCOMES ATTRIB TO ADV EVENT UPDATED. B5. DESCRIBE EVENT OR PROBLEM UPDATED. H6. PATIENT CODE AND IMPACT CODE UPDATED. IT WAS REPORTED THAT SHAFT BREAK OCCURRED AND RESULTED IN AN UNRETRIEVED DEVICE FRAGMENT. A 4.00 X 24 MM SYNERGY XD DRUG-ELUTING STENT WAS SELECTED FOR TREATMENT; HOWEVER, DURING THE PROCEDURE, A SHAFT BREAK OCCURRED. THE DEVICE WAS PARTIALLY EXPLANTED; HOWEVER, A PORTION REMAINS IMPLANTED. THE PROCEDURE WAS COMPLETED USING AN ALTERNATE DEVICE, AND NO PATIENT COMPLICATIONS WERE REPORTED. IT WAS FURTHER REPORTED THAT THE TARGET LESION WAS OVER 50% STENOSED, WITH A 90-DEGREE BEND LOCATED IN THE SEVERELY TORTUOUS AND SEVERELY CALCIFIED LEFT CIRCUMFLEX ARTERY (LCX). THE SHAFT BREAK OCCURRED AT THE DISTAL SECTION OF THE DEVICE AND A SNARE WAS USED IN AN UNSUCCESSFUL ATTEMPT TO REMOVE THE DEVICE FRAGMENT. A STENT WAS THEN PLACED OVER THE DEVICE FRAGMENT. THIS RESULTED IN A LCX DISSECTION, WHICH COULD NOT BE RETRIEVED OR REOPENED. H6. EVALUATION CONCLUSION CODES CORRECTED. IT WAS REPORTED THAT SHAFT BREAK OCCURRED AND RESULTED IN AN UNRETRIEVED DEVICE FRAGMENT. A 4.00 X 24 MM SYNERGY XD DRUG-ELUTING STENT WAS SELECTED FOR TREATMENT; HOWEVER, DURING THE PROCEDURE, A SHAFT BREAK OCCURRED. THE DEVICE WAS PARTIALLY EXPLANTED; HOWEVER, A PORTION REMAINS IMPLANTED. THE PROCEDURE WAS COMPLETED USING AN ALTERNATE DEVICE, AND NO PATIENT COMPLICATIONS WERE REPORTED. IT WAS FURTHER REPORTED THAT THE TARGET LESION WAS OVER 50% STENOSED, WITH A 90-DEGREE BEND LOCATED IN THE SEVERELY TORTUOUS AND SEVERELY CALCIFIED LEFT CIRCUMFLEX ARTERY (LCX). THE SHAFT BREAK OCCURRED AT THE DISTAL SECTION OF THE DEVICE AND A SNARE WAS USED IN AN UNSUCCESSFUL ATTEMPT TO REMOVE THE DEVICE FRAGMENT. A STENT WAS THEN PLACED OVER THE DEVICE FRAGMENT. THIS RESULTED IN A LCX DISSECTION, WHICH COULD NOT BE RETRIEVED OR REOPENED.
20744731 | 2024 | SYNERGY | Female | ['Difficult to Advance', 'Vascular Dissection', 'Reduced Blood Flow'] | IT WAS REPORTED THAT A DISSECTION OCCURRED. THE 80% STENOSED TARGET LESION WAS LOCATED IN THE MILDLY TORTUOUS AND MODERATELY CALCIFIED LEFT ANTERIOR DESCENDING ARTERY. FOLLOWING PRE-DILATATION, RESISTANCE WAS ENCOUNTERED, AND A 38 X 2.50 PROMUS PREMIER DRUG-ELUTING STENT WAS DEPLOYED AND POST DILATED. A PROXIMAL EDGE DISSECTION OCCURRED AND AN ADDITIONAL STENT WAS DEPLOYED TO COVER THE PROXIMAL EDGE DISSECTION. THE PROCEDURE WAS COMPLETED, AND NO FURTHER PATIENT COMPLICATIONS WERE REPORTED. B5. DESCRIBE EVENT OR PROBLEM UPDATED, D. SUSPECT MEDICAL DEVICE UPDATED. IT WAS REPORTED THAT A DISSECTION OCCURRED. THE 80% STENOSED TARGET LESION WAS LOCATED IN THE MILDLY TORTUOUS AND MODERATELY CALCIFIED LEFT ANTERIOR DESCENDING ARTERY. FOLLOWING PRE-DILATATION, RESISTANCE WAS ENCOUNTERED, AND A 38 X 2.50 PROMUS PREMIER DRUG-ELUTING STENT WAS DEPLOYED AND POST DILATED. A PROXIMAL EDGE DISSECTION OCCURRED, AND AN ADDITIONAL STENT WAS DEPLOYED TO COVER THE PROXIMAL EDGE DISSECTION. THE PROCEDURE WAS COMPLETED, AND NO FURTHER PATIENT COMPLICATIONS WERE REPORTED. IT WAS FURTHER REPORTED THAT A 2.25 X 16 MM SYNERGY DRUG-ELUTING STENT (DES) WAS IMPLANTED AND NOT A 2.50 X 38 MM PROMUS PREMIER DES AS PREVIOUSLY REPORTED. THE STENT WAS FULLY DEPLOYED AT 9 ATMOSPHERES WITHOUT ISSUES. ADDITIONALLY, THE DISSECTION WAS FLOW-LIMITING AND GRADED AS C. AFTER THE PROCEDURE, THE PATIENT WAS STABLE AND DISCHARGED.
20744867 | 2024 | XIENCE ALPINE | Male | ['Difficult to Remove', 'Difficult to Advance', 'Device Dislodged or Dislocated', 'Foreign Body In Patient'] | THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED ISSUE AND TREATMENT APPEARS TO BE RELATED TO OPERATIONAL CONTEXT. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. IT WAS REPORTED THE PROCEDURE WAS TO TREAT A MODERATELY CALCIFIED AND TORTUOUS LESION IN THE PROXIMAL LEFT ANTERIOR DESCENDING (LAD) ARTERY. THE LESION WAS PRE-DILATED WITH A SEMI COMPLAINT BALLOON. THE 3.50X18MM XIENCE ALPINE STENT DELIVERY SYSTEM (SDS) WAS ADVANCED WITH DIFFICULTY TO THE LESION WHEN IT WAS DETERMINED THE STENT WOULD NOT COVER THE LESION THEREFORE IT WAS DECIDED TO EXCHANGE FOR ANOTHER ONE. DURING REMOVAL THE SDS MET RESISTANCE WITH THE 6FR GUIDING CATHETER COULD NOT BE RETRACTED. IT WAS NOTED THE STENT HAD MOVED A FEW MILLIMETERS ON THE BALLOON THEREFORE IT WAS DECIDED TO DEPLOY THE STENT IN THE RADIAL ARTERY. ANOTHER XIENCE SIERRA WAS USED TO COMPLETE THE PROCEDURE. A CLINICALLY SIGNIFICANT DELAY WAS REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED.
20744932 | 2024 | SYNERGY | Male | ['Material Deformation', 'Vascular Dissection'] | IT WAS REPORTED THAT A DISSECTION OCCURRED. THE 80% STENOSED TARGET LESION WAS LOCATED IN THE MILDLY TORTUOUS AND MODERATELY CALCIFIED LEFT ANTERIOR DESCENDING ARTERY. THE LESION WAS PRE-DILATED AND A 3.50 X 20 MM SYNERGY DRUG-ELUTING STENT WAS ADVANCED WITH RESISTANCE AND DEPLOYED. FOLLOWING POST DILATION OF THE STENT, THE PHYSICIAN IDENTIFIED A PROXIMAL EDGE DISSECTION, WHICH REQUIRED A SMALLER SYNERGY STENT TO COVER THE DISSECTION. THE PROCEDURE WAS COMPLETED AND NO FURTHER COMPLICATIONS WERE REPORTED.
20745105 | 2024 | SYNERGY | Female | ['Difficult to Advance', 'Material Deformation', 'Vascular Dissection'] | IT WAS REPORTED THAT VESSEL DISSECTION OCCURRED. THE 70% STENOSED TARGET LESION WAS LOCATED IN THE MILDLY TORTUOUS AND MODERATELY CALCIFIED LEFT ANTERIOR DESCENDING ARTERY. THE LESION WAS PREDILATED USING A 2.50 X 12 MM SEMI-COMPLIANT BALLOON. A 2.50 X 32 MM SYNERGY II DRUG-ELUTING STENT WAS ADVANCED WITH RESISTANCE FOR TREATMENT. HOWEVER, AFTER DEPLOYING AND POST DILATING THE STENT, A PROXIMAL EDGE DISSECTION WAS OBSERVED IN FLUOROSCOPY AND COVERED USING ANOTHER SMALLER SYNERGY STENT. STENT DAMAGE WAS ALSO NOTED. THE PROCEDURE WAS COMPLETED WITH NO FURTHER PATIENT COMPLICATIONS.
20745920 | 2024 | SYNERGY | Female | ['Material Deformation', 'Vascular Dissection'] | IT WAS REPORTED THAT A DISSECTION OCCURRED. THE 80% STENOSED TARGET LESION WAS LOCATED IN THE MILDLY TORTUOUS AND MODERATELY CALCIFIED LEFT CIRCUMFLEX ARTERY (LCX). AFTER PREDILATION OF THE LESION, A 2.25 X 28 SYNERGY STENT WAS ADVANCED WITH RESISTANCE, DEPLOYED AND POST DILATED. A PROXIMAL STENT EDGE DISSECTION WAS OBSERVED. THE DISSECTION WAS COVERED WITH A SMALLER SYNERGY STENT AND THE PROCEDURE WAS COMPLETED WITHOUT ANY FURTHER PATIENT COMPLICATIONS. STENT DAMAGE WAS ALSO NOTED.
20746705 | 2024 | SYNERGY XD | Unknown | ['Defective Device', 'Thrombosis/Thrombus', 'Reduced Blood Flow'] | IT WAS REPORTED THAT STENT THROMBOSIS OCCURRED. A 3.00 X 24 SYNERGY DRUG-ELUTING STENT WAS PLACED FOR TREATMENT. HOWEVER, POST DEPLOYMENT, SLOW FLOW, RECOILING OF PROXIMAL SITE PART OF THE STENT AND STENT THROMBOSIS WERE NOTED. THE PROCEDURE WAS COMPLETED, AND NO FURTHER PATIENT COMPLICATIONS WERE REPORTED, AND THE PATIENT WAS EXPECTED TO FULLY RECOVER. IT WAS REPORTED THAT STENT THROMBOSIS OCCURRED. A 3.00 X 24 SYNERGY DRUG-ELUTING STENT WAS PLACED FOR TREATMENT. HOWEVER, POST DEPLOYMENT, SLOW FLOW, RECOILING OF PROXIMAL SITE PART OF THE STENT AND STENT THROMBOSIS WERE NOTED. THE PROCEDURE WAS COMPLETED, AND NO FURTHER PATIENT COMPLICATIONS WERE REPORTED, AND THE PATIENT WAS EXPECTED TO FULLY RECOVER. IT WAS FURTHER REPORTED THAT THE PATIENT PRESENTED WITH ST-ELEVATION MYOCARDIAL INFARCTION (STEMI) AND CHEST PAIN. THE 100% STENOSED TARGET LESION WAS LOCATED IN THE NON-TORTUOUS AND NON-CALCIFIED LEFT ANTERIOR DESCENDING ARTERY (LAD). THE THROMBOSIS, DISCOVERED DURING THE PERCUTANEOUS TRANSLUMINAL CORONARY ANGIOPLASTY (PTCA) VIA ANGIOGRAM, WAS CAUSED BY RECOILING OF THE STENT. THE PATIENT WAS ON ANTIPLATELET THERAPY AT THE TIME OF PROCEDURE, AND TREATMENT WITH PLAIN OLD BALLOON ANGIOPLASTY (POBA) WAS PERFORMED TO TREAT THE THROMBOSIS.
20755974 | 2024 | ONYX FRONTIER | Male | ['Device Dislodged or Dislocated', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING AN ONYX FRONTIER CORONARY DRUG ELUTING STENT, STENT DISLODGEMENT OCCURRED WHEN PULLING BACK FROM THE DISTAL TO MID PART OF THE LESION. THE LESION WAS SEVERELY TORTUOSITY, SEVERELY CALCIFICATION WITH A 90% STENOSIS AND LOCATED IN THE MID-DISTAL RIGHT CORONARY ARTERY (RCA). THE LESION WAS PRE-DILATED. THE DEVICE DID NOT PASS THROUGH A PREVIOUSLYDEPLOYED STENT. RESISTANCE WAS NOT ENCOUNTERED WHEN ADVANCING THE DEVICE. EXCESSIVE FORCE WAS NOT USED. THE STENT WAS SUCCESSFULLY REMOVED USING BALLOON INFLATION. THE PATIENT IS ALIVE WITH NO FURTHER INJURY. ADDITIONAL INFORMATION: IT WAS DETAILED THAT THE STENT WAS ADVANCED PAST THE LESION AND THEN WAS PULLED BACK FOR PROPER OVERLAP TO THE DISTAL STENT THAT IT WAS PREVIOUSLY PLACED WITH NO ISSUES. RESISTANCE WAS EXPERIENCED IN PULLING THE STENT BACK WHICH WAS BELIEVED TO HAVE THE STENT DISLODGED. EXCESSIVE FORCE WAS USED. THE STENT WAS REMOVED DURING THE PROCEDURE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: THE PREVIOUSLY PLACED DISTAL STENT WAS AN ONYX FRONTIER STENT. THIS ONYX FRONTIER STENT WAS NOT FULLY EXPANDED DUE TO RIGID CALCIUM. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: ANNEX B CODE. ANNEX D CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: THE PREVIOUSLY DEPLOYED DISTAL ONYX FRONTIER STENT WAS NOT FULLY EXPANDED DUE TO RIGID CALCIUM, AND THE NEW 3.5X26MM ONYX FRONTIER WAS BEING ATTEMPTED TO BE DEPLOYED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20756712 | 2024 | ONYX TRUCOR | Male | ['Improper or Incorrect Procedure or Method', 'Deformation Due to Compressive Stress', 'Detachment of Device or Device Component', 'Device Dislodged or Dislocated', 'Material Deformation', 'Positioning Problem', 'Insufficient Information'] | IT WAS REPORTED THAT DURING AN INTERVENTIONAL RADIOLOGY PROCEDURE INVOLVING ONE ONYX TRUCOR CORONARY DRUG ELUTING STENT, THE STENT COULD NOT PASS THROUGH THE LESION. THE STENT WAS ALSO REPORTED TO BE DEFORMED IN VIVO DURING POSITIONING/ADVANCEMENT. THE STENT WAS STRIPPED, AND THE SHAFT BROKE WHILE BEING RETRIEVED. THE DEVICE WAS BEING USED IN A SEVERELY CALCIFIED, MODERATELY TORTUOUS LESION IN THE MID MIDDLE CEREBRAL ARTERY (MCA). THE DEVICE HAD BEEN INSPECTED BEFORE USE WITH NO ISSUES. NEGATIVE PREP WAS PERFORMED WITH NO ISSUES. THE LESION WAS NOT PRE-DILATED. THE DEVICE DID NOT PASS THROUGH A PREVIOUSLY DEPLOYED STENT. RESISTANCE WAS ENCOUNTERED WHEN ADVANCING THE DEVICE. THE PROCEDURE WAS COMPLETED USING A DIFFERENT BRAND OF STENT. IT IS BELIEVED THAT THE EVENT WAS DUE TO USE OF THE DEVICE IN DIFFICULT LESION MORPHOLOGY/ANATOMY OR PROCEDURAL RELATED. THE PATIENT IS ALIVE WITH NO FURTHER INJURY. PLEASE NOTE THAT THIS DEVICE (ONYX TRUCOR) IS NOT MARKETED IN THE UNITED STATES; HOWEVER, THE DEVICE IS DEEMED THE SAME AS THE UNITED STATES MARKETED DEVICE (RESOLUTE ONYX RX). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: THE SHAFT BROKE INTO TWO SEPARATE PIECES WHILE BEING RETRIEVED. THE DEVICE BROKE INTO TWO SEPARATE PIECES OUTSIDE OF THE PATIENT SO IT WAS REMOVED AND DOES NOT REMAIN IN THE PATIENT. THE STENT SCRAPED OFF/DISLODGED INSIDE THE CATHETER, WAS REMOVED AND DOES NOT REMAIN INSIDE THE PATIENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. PRODUCT ANALYSIS: THE DEVICE RETURNED WITH A DETACHMENT ON THE HYPOTUBE. KINKS WERE EVIDENT ON THE HYPOTUBE. THE HYPOTUBE MATERIAL WAS OVAL AND JAGGED ON BOTH SIDES OF THE DETACHMENT SITE. THE STENT WAS NOT PRESENT ON THE BALLOON AND DID NOT RETURN ALONGSIDE THE DEVICE. BALLOON FOLDS WERE INTACT ON DEVICE RETURN. CRIMP IMPRESSIONS VISIBLE ON EXPOSED BALLOON SURFACE. DEFORMATION WAS EVIDENT TO DISTAL TIP. NO OTHER DAMAGE EVIDENT TO THE REMAINDER OF THE DEVICE. IMAGE ANALYSIS: MULTIPLE IMAGES WERE PROVIDED SHOWING THE CEREBRAL VASCULATURE. A TIGHT STENOTIC LESION IS PRESENT IN THE MIDDLE CEREBRAL ARTERY. NO TREATMENT OF THIS LESION WAS SHOWN IN THE IMAGES ALTHOUGH IT WAS REPORTED THAT THE LESION TREATMENT WAS COMPLETED USING A DIFFERENT BRAND OF STENT. FROM THE IMAGE IT CAN BE CONFIRMED THAT THE INTERNAL CAROTID ARTERY WAS SEVERELY TORTUOUS AND THERE IS NO EVIDENCE TO SHOW THAT THE TRUCOR STENT WAS NOT ABLE TO BE DELIVERED. THE IMAGES SUGGEST THAT THE SEVERE TORTUOSITY MAY HAVE CONTRIBUTED TO THE FAILURE TO DELIVER THE STENT. THE IMAGES DO NOT PROVIDE EVIDENCE TO CONFIRM THE DISLODGEMENT OR SHAFT BREAKAGE. ADDITIONAL INFORMATION: THE STENT SCRAPED OFF/DISLODGED INSIDE THE NON-MEDTRONIC CATHETER, WAS REMOVED AND DOES NOT REMAIN INSIDE THE PATIENT. IT WAS NOT BELIEVED THAT THE CATHETER HAD CAUSED OR CONTRIBUTED TO THE STENT DISLODGING. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: ANNEX D CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20765408 | 2024 | SYNERGY XD | Unknown | ['Break', 'Difficult to Remove', 'Detachment of Device or Device Component', 'Foreign Body In Patient'] | IT WAS REPORTED THAT CATHETER ENTRAPMENT, REMOVAL DIFFICULTY, SHAFT BREAK AND UNRETRIEVED DEVICE FRAGMENT OCCURRED. THE 99% STENOSED TARGET LESION WAS LOCATED IN THE SEVERELY TORTUOUS AND SEVERELY CALCIFIED RIGHT CORONARY ARTERY (RCA). A 3.50 X 48 MM SYNERGY XD DRUG-ELUTING STENT (DES) WAS DEPLOYED. HOWEVER, POST STENT DEPLOYMENT, THE STENT BALLOON COULD NOT BE REMOVED. THE SHAFT OF THE STENT DELIVERY SYSTEM (SDS) FRACTURED, LEAVING THE STENT BALLOON STUCK IN THE PROXIMAL AREA OF THE DES. THE PHYSICIAN THEN ATTEMPTED TO USE A SNARE TO REMOVE THE DEVICE FRAGMENT, BUT IT WAS UNSUCCESSFUL. THE PATIENT WAS SENT TO SURGERY, AND NO FURTHER ISSUES WERE REPORTED. DEVICE EVALUATED BY MFR: SYNERGY XD MR OUS 3.50 X 48MM WAS RETURNED FOR ANALYSIS. A VISUAL AND TACTILE INSPECTION OF THE HYPOTUBE PROFILE IDENTIFIED A BREAK AT THE EXCHANGE PORT. THE REMAINING SECTION, WITH THE MANIFOLD ATTACHED, DID NOT RETURN FOR ANALYSIS. A VISUAL, TACTILE AND MICROSCOPIC INSPECTION OF THE SHAFT POLYMER EXTRUSION IDENTIFIED NO ISSUES WITH THE OUTER / MID-SHAFT SECTIONS OR THE INNER LUMEN OF THE DEVICE. THE DISTAL SECTION OF THE DEVICE RETURNED LOADED ON A GUIDEWIRE. THE POLYMER EXTRUSION WAS BUNCHED AND STRETCHED. A MICROSCOPIC ANALYSIS OF THE BALLOON PROFILED REVEALED THAT IT APPEARED THE BALLOON HAD BEEN SUBJECTED TO POSITIVE PRESSURE HOWEVER, AN EXAMINATION OF THE BALLOON MATERIAL FOUND NO DEFECTS. A MICROSCOPIC EXAMINATION OF THE PROXIMAL AND DISTAL MARKERBANDS IDENTIFIED NO DAMAGE. THE BUMPER TIP SHOWED SIGNS OF DISTAL TIP DAMAGE (FLARED). A DIMENSIONAL TESTING IDENTIFIED THERE WAS NO STENT RETURNED THEREFORE IT WAS NOT POSSIBLE TO TAKE A READING OF THE CRIMPED STENT OD (OUTER DIAMETER). THE MAX STENT PROFILE AT THE TIME OF MANUFACTURING WAS NOTED TO BE WITHIN MAX CRIMPED STENT PROFILE MEASUREMENT. THE EXACT MAX STENT PROFILE CANNOT BE DETERMINED DUE TO THE NO RETURN OF THE MANIFOLD WHICH HAS THE BATCH NUMBER ATTACHED TO IT HOWEVER THE MAX STENT PROFILE OF THE DATA PULLED DID NOT EXCEED THE SPECIFICATION. IT WAS REPORTED THAT CATHETER ENTRAPMENT, REMOVAL DIFFICULTY, SHAFT BREAK AND UNRETRIEVED DEVICE FRAGMENT OCCURRED. THE 99% STENOSED TARGET LESION WAS LOCATED IN THE SEVERELY TORTUOUS AND SEVERELY CALCIFIED RIGHT CORONARY ARTERY (RCA). A 3.50 X 48 MM SYNERGY XD DRUG-ELUTING STENT (DES) WAS DEPLOYED. HOWEVER, POST STENT DEPLOYMENT, THE STENT BALLOON COULD NOT BE REMOVED. THE SHAFT OF THE STENT DELIVERY SYSTEM (SDS) FRACTURED, LEAVING THE STENT BALLOON STUCK IN THE PROXIMAL AREA OF THE DES. THE PHYSICIAN THEN ATTEMPTED TO USE A SNARE TO REMOVE THE DEVICE FRAGMENT, BUT IT WAS UNSUCCESSFUL. THE PATIENT WAS SENT TO SURGERY, AND NO FURTHER ISSUES WERE REPORTED.
20772417 | 2024 | ORSIRO MISSION 3.0/13 | Unknown | ['Patient Device Interaction Problem', 'Thrombosis/Thrombus'] | COMBINATION PRODUCT: YES. THE ORSIRO MISSION DRUG-ELUTING STENT SYSTEM WAS SELECTED FOR TREATMENT OF A MILDLY CALCIFIED, TOTAL OCCLUDED, LESION IN THE MILDLY TORTUOUS PART OF THE MID LAD. AFTER PRE-DILATATION, 2 ORSIRO MISSION OVERLAPPING STENTS WERE IMPLANTED AND GOOD RESULTS WERE OBTAINED WITH IVUS. IMMEDIATELY AFTERWARDS, INCIPIENT THROMBOSIS OF THE STENTS WAS OBSERVED. MEDICATION WAS GIVEN AND POST-DILATATION WAS PERFORMED WITH GOOD RESULTS. IN LESS THAN 24 HRS PATIENT PRESENTED AGAIN WITH NEW THROMBOSIS WHICH WAS TREATED AGAIN BY POST-DILATATION WITH A NON-COMPLIANT BALLOON. GOOD EXPANSION AND FREE EDGES OF THE LESION WERE OBSERVED. NEW THROMBOSIS OCCURRED 4 DAYS LATER. THIS IS 2 OF 2. COMBINATION PRODUCT: YES. THE AFFECTED PRODUCT WAS NOT RETURNED TO BIOTRONIK AND COULD THEREFORE, NOT BE SUBJECTED TO A TECHNICAL INVESTIGATION. IMAGES OF THE CASE SUCH AS ANGIOGRAPHIES, IVUS OR OCT COULD NOT BE OBTAINED. REVIEW OF THE PRODUCTION DOCUMENTATION CONFIRMED THAT THE DEVICE IN QUESTION WAS MANUFACTURED ACCORDING TO SPECIFICATIONS AND FULFILLED ALL THE REQUIREMENTS OF IN? PROCESS AND FINAL INSPECTION. BASED ON THE REVIEWED DOCUMENTATION NO DEVICE DEFICIENCY OR MANUFACTURING-RELATED ROOT CAUSE COULD BE IDENTIFIED.
20772466 | 2024 | ORSIRO MISSION 2.5/35 | Unknown | ['Patient Device Interaction Problem', 'Thrombosis/Thrombus'] | THE ORSIRO MISSION DRUG-ELUTING STENT SYSTEM WAS SELECTED FOR TREATMENT OF A MILDLY CALCIFIED, TOTAL OCCLUDED, LESION IN THE MILDLY TORTUOUS PART OF THE MID LAD. AFTER PRE-DILATATION, 2 ORSIRO MISSION OVERLAPPING STENTS WERE IMPLANTED AND GOOD RESULTS WERE OBTAINED WITH IVUS. IMMEDIATELY AFTERWARDS, INCIPIENT THROMBOSIS OF THE STENTS WAS OBSERVED. MEDICATION WAS GIVEN AND POST-DILATATION WAS PERFORMED WITH GOOD RESULTS. IN LESS THAN 24 H PATIENT PRESENTED AGAIN WITH NEW THROMBOSIS WHICH WAS TREATED AGAIN BY POST-DILATATION WITH A NON-COMPLIANT BALLOON. GOOD EXPANSION AND FREE EDGES OF THE LESION WERE OBSERVED. NEW THROMBOSIS OCCURRED 4 DAYS LATER. THIS IS 1 OF 2. COMBINATION PRODUCT: YES. COMBINATION PRODUCT: YES, THE AFFECTED PRODUCT WAS NOT RETURNED TO BIOTRONIK AND COULD THEREFORE NOT BE SUBJECTED TO A TECHNICAL INVESTIGATION. IMAGES OF THE CASE SUCH AS ANGIOGRAPHIES, IVUS OR OCT COULD NOT BE OBTAINED. REVIEW OF THE PRODUCTION DOCUMENTATION CONFIRMED THAT THE DEVICE IN QUESTION WAS MANUFACTURED ACCORDING TO SPECIFICATIONS AND FULFILLED ALL THE REQUIREMENTS OF IN?PROCESS AND FINAL INSPECTION. BASED ON THE REVIEWED DOCUMENTATION NO DEVICE DEFICIENCY OR MANUFACTURING-RELATED ROOT CAUSE COULD BE IDENTIFIED.
20777932 | 2024 | PROMUS ELITE | Male | ['Fracture', 'Insufficient Information'] | IT WAS REPORTED THAT STENT FRACTURE OCCURRED. THE PATIENT PRESENTED WITH ST ELEVATION MYOCARDIAL INFARCTION AND WAS BEING TREATED WITH PERCUTANEOUS CORONARY INTERVENTION (PCI) OF A TARGET LESION WAS LOCATED IN THE LEFT ANTERIOR DESCENDING (LAD) ARTERY. FOLLOWING PRE-DILATION OF THE MINOR CALCIFICATION WITH A SEMI-COMPLIANT BALLOON, A 28 X 3.00 PROMUS ELITE DRUG-ELUTING STENT WAS DEPLOYED. DURING POST-DILATION WITH A NON-COMPLIANT BALLOON, THE STENT FRACTURED. THE PROCEDURE WAS COMPLETED WITH A DIFFERENT DEVICE. THERE WERE NO PATIENT COMPLICATIONS REPORTED AND THE PATIENT WAS NOTED TO BE STABLE POST-PROCEDURE.
20792169 | 2024 | SYNERGY MEGATRON | Unknown | ['Device Dislodged or Dislocated', 'No Clinical Signs, Symptoms or Conditions'] | IT WAS REPORTED THAT STENT DISLODGMENT OCCURRED. THE 4.00 X 16MM SYNERGY MEGATRON MR DRUG-ELUTING STENT WAS SELECTED FOR USE. DURING THE PROCEDURE THE STENT DISLODGED IN THE PROXIMAL RIGHT CORONARY ARTERY AND WAS REMOVED WITH A SNARE. THE PROCEDURE WAS COMPLETED WITH ANOTHER OF THE SAME DEVICE. THERE WERE NO PATIENT COMPLICATIONS.
20798261 | 2024 | SYNERGY XD | Unknown | ['Device-Device Incompatibility', 'Migration', 'No Clinical Signs, Symptoms or Conditions'] | IT WAS REPORTED THAT STENT MIGRATION OCCURRED WHICH REQUIRED INTERVENTION. THE TARGET LESION WAS LOCATED IN THE RIGHT CORONARY ARTERY. FOLLOWING A PRE-INTRAVASCULAR ULTRASOUND (IVUS) ON THE LESION, A 5.00 X 16MM SYNERGY XD DRUG-ELUTING STENT WAS DEPLOYED FOR TREATMENT. HOWEVER, WHEN THE PHYSICIAN BALLOONED THE STENT WITH A NON-COMPLIANT BALLOON, IT WAS OBSERVED THAT THE STENT MOVED FROM THE ORIGINAL POSITION WITHIN THE VESSEL. SUBSEQUENTLY, THE STENT WAS POST DILATED AND DEPLOYED ANOTHER SYNERGY STENT MORE PROXIMALLY TO COVER THE AREA THAT WAS MISSED. THE PROCEDURE WAS COMPLETED SUCCESSFULLY WITH NO PATIENT COMPLICATIONS REPORTED.
20800862 | 2024 | SYNERGY XD | Unknown | ['Break', 'Failure to Advance', 'Device-Device Incompatibility', 'No Clinical Signs, Symptoms or Conditions'] | IT WAS REPORTED THAT SHAFT BREAK OCCURRED, REQUIRING ADDITIONAL INTERVENTION. THE PATIENT PRESENTED WITH DISEASE THROUGHOUT THE WHOLE VESSEL VIA ANGIOGRAPHY AND WAS UNDERGOING PERCUTANEOUS CORONARY INTERVENTION. THE TARGET LESION WAS LOCATED IN THE MODERATELY CALCIFIED RIGHT CORONARY ARTERY. A 2.50 X 32MM SYNERGY XD DRUG-ELUTING STENT WAS ADVANCED FOR TREATMENT. THE FIRST STENT PLACEMENT WAS SUCCESSFUL WITH THE USE OF A NON-BOSTON SCIENTIFIC (BSC) GUIDE EXTENSION CATHETER TO AID IN DELIVERABILITY AND SUPPORT. HOWEVER, WHEN THE PHYSICIAN ATTEMPTED TO ADVANCE THE SECOND STENT IN BETWEEN A NON-BOSTON SCIENTIFIC (BSC) GUIDE EXTENSION CATHETER AND THE FIRST STENT, THE STENT GOT CAUGHT ON THE PROXIMAL EDGE OF THE FIRST STENT AND POSSIBLY ON CALCIUM AND RESISTANCE WAS ENCOUNTERED. CONSEQUENTLY, THE BALLOON CATHETER BROKE OFF DUE TO THE RESISTANCE INSIDE THE VESSEL, LEAVING THE BALLOON AND STENT SEPARATED FROM THE DELIVERY SYSTEM. THE PHYSICIAN WAS ABLE TO RETRIEVE THE STENT AND THE DISTAL BALLOON CATHETER WITH A SNARE. A NEW STENT WAS PLACED, AND THE PROCEDURE WAS COMPLETED SUCCESSFULLY. NO PATIENT COMPLICATIONS WERE REPORTED.
20809078 | 2024 | ONYX FRONTIER | Unknown | ['Device Dislodged or Dislocated', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING ONE ONYX FRONTIER CORONARY DRUG ELUTING STENT, THE STENT UNINTENTIONALLY DISLODGED FROM THE STENT DELIVERY SYSTEM (SDS) BALLOON, IN AN EXTREMELY TORTUOUS VESSEL, DURING DELIVERY TO THE LESION. THE LESION WAS SEVERELY TORTUOUS, AND SEVERELY CALCIFIED AND LOCATED IN THE MID-DISTAL RIGHT CORONARY ARTERY (RCA). THE DEVICE WAS INSPECTED BEFORE USE WITH NO ISSUES. NEGATIVE PREP WAS PERFORMED WITH NO ISSUES. RESISTANCE WAS ENCOUNTERED WHEN ADVANCING THE DEVICE. THE DISLODGED STENT WAS NOT REMOVED. THE DISLODGED STENT STAYED ON WIRE, AND ONE 2X12MM SPRINTER BALLOON AND ONE 3X15MM EUPHORA BALLOON WERE USED TO DILATE/EXPAND AND DEPLOY THE STENT IN THE VESSEL. THE PATIENT RECEIVED ALMOST 10GY IN DOSE. PREVIOUS TO THIS EVENT A NO CROSS STENT WAS DAMAGED ON INSPECTING UPON REMOVAL FROM THE BODY. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION: THE LESION WAS PRE-DILATED. THE PATIENT RECEIVED ALMOST 10GY IN DOSE OF RADIATION. PREVIOUS TO THIS EVENT A NO CROSS MEDTRONIC STENT WAS DAMAGED ON INSPECTING UPON REMOVAL FROM THE BODY. THERE IS NO COMPLAINT AGAINST THE SPRINTER AND EUPHORA BALLOONS USED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: THE MEDTRONIC NO CROSS STENT WHICH WAS DAMAGED ON INSPECTION AFTER REMOVAL WAS INSPECTED BEFORE USE WITH NO ISSUES NOTED. RESISTANCE WAS NOTED WHILE ADVANCING THIS DEVICE TO THE LESION. THE MEDTRONIC NO CROSS STENT WAS NOT IMPLANTED IN THE PATIENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20809823 | 2024 | ONYX FRONTIER | Male | ['Device Dislodged or Dislocated', 'Chest Pain', 'Non specific EKG/ECG Changes', 'Obstruction/Occlusion', 'Vascular Dissection', 'Device Embedded In Tissue or Plaque', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. THE PATIENT PRESENTED WITH SUDDEN ONSET OF CHEST PAIN AND SHORTNESS OF BREATH. IT WAS REPORTED THAT DURING A PROCEDURE AN ATTEMPT WAS MADE TO USE ONE ONYX FRONTIER CORONARY DRUG ELUTING STENT TO TREAT LESIONS IN THE LEFT ANTERIOR DESCENDING ARTERY (LAD) AND DISTAL RCA AND STENT DISLODGEMENT OCCURRED. IT WAS DETAILED THAT THE STENT WAS MISSING AND THE PHYSICIAN TRIED A WIRE TECHNIQUE TO FIND AND REMOVE THE STENT. THE STENT FLOATED TO THE PROXIMAL RIGHT CORONARY ARTERY (RCA) AND DUE TO MULTIPLE WIRES USED DURING THE PROCEDURE A SPIRAL CORONARY DISSECTION OCCURRED. THE STENT WAS ABANDONED. THE PATIENT ENDED UP HAVING 6 STENTS IMPLANTED IN THE WHOLE RCA. A FALSE LUMEN WAS INTENTIONALLY STENTED UP TO THE MID VESSEL WHERE IT WAS POSSIBLE TO RE ENTER INTO THE TRUE LUMEN AND CONNECT TRUE LUMEN TO TRUE LUMEN DISTALLY. THE SPIRAL DISSECTION DID RESULT IN LOSS OF AN ACUTE MARGINAL LATERAL BRANCH, WHICH RESULTED IN SOME LIKELY CHEST DISCOMFORT AND ABNORMAL EKG. NO FURTHER PATIENT INJURY OCCURRED. ADDITIONAL INFORMATION: EVENT DATE UPDATED. SECTION A. PATIENT INFORMATION UPDATED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20810187 | 2024 | ONYX TRUSTAR | Male | ['Burst Container or Vessel', 'Leak/Splash', 'Difficult to Remove', 'Insufficient Information'] | PLEASE NOTE THAT THIS DEVICE (ONYX TRUSTAR) IS NOT MARKETED IN THE UNITED STATES; HOWEVER, THE DEVICE IS DEEMED THE SAME AS THE UNITED STATES-MARKETED DEVICE (ONYX FRONTIER). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING ONE ONYX TRUSTAR CORONARY DRUG ELUTING STENT, THERE WAS A BALLOON BURST OR LEAK DURING STENT DEPLOYMENT AT NOMINAL PRESSURE. THE STENT COULD NOT BE REMOVED FROM THE PATIENT AND HAD TO BE INFLATED IN THE ARTERY RADIALIS. THE DEVICE WAS INSPECTED WITH NO ISSUES. NEGATIVE PREP WAS PERFORMED WITH NO ISSUES. THE LESION WAS PRE-DILATED. THE PATIENT HAS RECOVERED FULLY. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION: THE LESION BEING TREATED WAS NON-TORTUOUS, HAD 30% CALCIFICATION, 70% STENOSIS AND LOCATED IN THE OSTIAL LEFT MAIN CORONARY ARTERY. IT WAS LATER CLARIFIED THAT A LEAK OCCURRED FROM THE VERY BEGINNING WHEN ATTEMPTING TO INFLATE THE BALLOON. AN INFLATION PRESSURE OF 0 ATM WAS NOTED AS IT WAS NOT POSSIBLE TO INFLATE THE BALLOON DUE TO THE LEAK. THE ISSUE OCCURRED ON THE FIRST ATTEMPTED INFLATION. THE LESION WAS PRE-DILATED WITH A 3.0X8MM CUTTING BALLOON TWICE WITH GOOD RESULTS. THE DEVICE DID NOT PASS THROUGH A PREVIOUSLY DEPLOYED STENT. RESISTANCE WAS NOT ENCOUNTERED WHEN ADVANCING THE DEVICE, OR DURING THE WITHDRAWAL OF THE DEVICE. THE DEVICE WAS NOT MOVED OR REPOSITIONED IN THE LESION WHILE INFLATED. SECTIONS A. PATIENT INFORMATION AND B7.RELEVANT HISTORY UPDATED. INITIAL REPORTER PHONE NUMBER PROVIDED. DATE OF EVENT UPDATED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20831691 | 2024 | RESOLUTE INTEGRITY RX | Male | ['Difficult to Remove', 'Device Dislodged or Dislocated', 'Device Embedded In Tissue or Plaque', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. IT WAS REPORTED THAT ONE RESOLUTE INTEGRITY CORONARY DRUG ELUTING STENT DISLODGED DURING SURGERY AND COULD NOT BE REMOVED. THE DISLODGED STENT WAS COVERED WITH ANOTHER STENT AND ATTACHED TO THE WALL WITHOUT CAUSING ANY ADVERSE EFFECTS. THERE WAS NO FURTHER PATIENT INJURY REPORTED. ADDITIONAL INFORMATION: ANNEX D CODE. CORRECTION: ANNEX A CODE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20832524 | 2024 | XIENCE ALPINE | Unknown | ['Entrapment of Device', 'Material Rupture', 'Improper or Incorrect Procedure or Method', 'Migration', 'No Clinical Signs, Symptoms or Conditions'] | THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. IT WAS REPORTED THAT A STENT WAS IMPLANTED IN THE PROXIMAL LESION, A SECOND STENT WAS ATTEMPTED TO BE DEPLOYED IN THE DISTAL. IT SHOULD BE NOTED THAT THE INSTRUCTIONS FOR USE (IFU) STATES: WHEN TREATING MULTIPLE LESIONS WITHIN THE SAME VESSEL, STENT THE DISTAL LESION PRIOR TO STENTING THE PROXIMAL LESION. THE INVESTIGATION DETERMINED THE REPORTED DIFFICULTIES APPEAR TO BE RELATED TO USER ERROR. IT WAS REPORTED THAT A STENT WAS IMPLANTED IN THE PROXIMAL LESION, FOLLOWED BY THE REPORTED STENT IN THE DISTAL. IN THIS CASE, IT IS LIKELY THAT WHEN ATTEMPTING TO PASS THE REPORTED STENT THROUGH THE IMPLANTED STENT, AN INTERACTION OCCURRED. AN INTERACTION BETWEEN THE STENTS COULD CONTRIBUTE THE REPORTED MIGRATION AND ENTRAPMENT. ADDITIONALLY, INTERACTION BETWEEN THE IMPLANTED STENT AND THE REPORTED SDS COULD CONTRIBUTE TO BALLOON DAMAGE, RESULTING IN A MATERIAL RUPTURE. IT SHOULD BE NOTED THAT THE INSTRUCTIONS FOR USE STATE Â¿WHEN TREATING MULTIPLE LESIONS WITHIN THE SAME VESSEL, STENT THE DISTAL LESION PRIOR TO STENTING THE PROXIMAL LESION.Â¿ BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. H6 - MEDICAL DEVICE PROBLEM CODE 2017 CLARIFIER: DISTAL TO STENT. IT WAS REPORTED THAT THE PROCEDURE WAS PERFORMED TO TREAT A LESION IN THE LEFT CIRCUMFLEX (CX) CORONARY ARTERY WITH MILD CALCIFICATION, MILD TORTUOSITY AND 80% STENOSIS. EACH XIENCE ALPINE STENT WAS PREPPED PER INSTRUCTIONS OF USE. THE FIRST STENT, A 2.5MM X 38MM XIENCE ALPINE STENT WAS IMPLANTED IN THE PROXIMAL CX. THE SECOND STENT, A 2.5X18MM XIENCE ALPINE STENT, WAS ATTEMPTED TO BE DEPLOYED IN THE DISTAL CX. THE BALLOON WAS INFLATED WAS INFLATED ONCE TO 8 ATMOSPHERES AND A RUPTURE OCCURRED, CAUSING THE STENT TO BE DEPLOYED UNDER-EXPANDED. DILATATION WAS ATTEMPTED WITH A 2.5X12MM NC TREK BALLOON DILATION CATHETER, YET THE STENT COULD NOT BE LOCATED UNDER FLUOROSCOPY. THEREFORE, A 2.5X23MM XIENCE ALPINE STENT WAS ADVANCED TO BE DEPLOYED IN THE DISTAL CX. SUBSEQUENTLY, THE MISSING 2.5X18MM XIENCE STENT WAS LOCATED, DISLODGED IN THE STRUTS OF THE FIRST IMPLANTED STENT. THEREFORE, THE 2.5MM X 23MM XIENCE STENT WAS DEPLOYED TO PUSH THE DISLODGED STENT OUT DISTALLY AND DILATATION WAS PERFORMED ON BOTH STENTS IN THE DISTAL CX. THERE WAS NO ADVERSE PATIENT SEQUELA. NO ADDITIONAL INFORMATION WAS PROVIDED.
20851398 | 2024 | RESOLUTE INTEGRITY RX | Male | ['Material Too Rigid or Stiff', 'Bradycardia', 'Hemorrhage/Blood Loss/Bleeding', 'Low Blood Pressure/ Hypotension', 'Perforation', 'Rupture', 'Cardiac Tamponade', 'Loss of consciousness', 'Obstruction/Occlusion', 'Insufficient Information'] | IT WAS REPORTED THAT A CORONARY ARTERY RUPTURE AND PERFORATION OCCURRED AFTER IMPLANTATION OF A RESOLUTE INTEGRITY DRUG ELUTING STENT. THE PATIENT WAS UNDERGOING CORONARY ANGIOGRAPHY WHICH SHOWED THAT THE CIRCUMFLEX BRANCH (CX) WAS COMPLETELY OCCLUDED. A 3.5 GUIDE CATHETER WAS USED, AND ONE NON-MEDTRONIC GUIDEWIRE WAS PLACED THROUGH THE CATHETER TO THE OCCLUDED LESION OF THE CX. AFTER REPEATED ATTEMPTS, THE GUIDEWIRE FAILED TO PASS THROUGH THE OCCLUDED LESION. UNDER THE SUPPORT OF A MICROCATHETER, THE GUIDEWIRE WAS CHANGED TO PASS THROUGH THE OCCLUDED LESION. THEN A NON-MEDTRONIC BALLOON WAS USED TO DILATE THE OCCLUDED LESION, AND THE 2.25X18MM RESOLUTE INTEGRITY STENT WAS IMPLANTED AT THE LESION. AFTER THE STENT WAS DILATED, THE ANGIOGRAPHY SHOWED EXTRAVASATION OF CONTRAST AGENT, AND IT WAS CONSIDERED THAT THE CORONARY ARTERY RUPTURED AND PERFORATED. BLOOD FLOWED INTO THE PERICARDIUM. A TOTAL OF THREE STENTS WERE IMPLANTED AND THE SURGERY WAS SUCCESSFUL. AFTER THE SURGERY WAS COMPLETED, THE PATIENT RETURNED TO THE WARD SAFELY. APPROXIMATELY TEN MINUTES POST-SURGERY THE PATIENT SUDDENLY LOST CONSCIOUSNESS, DID NOT RESPOND TO CALLS, AND THEIR HEART RATE AND BLOOD PRESSURE SLOWED DOWN. CARDIAC TAMPONADE WAS CONSIDERED. PERICARDIOCENTESIS WAS IMMEDIATELY PERFORMED, DRAINED 300ML OF BLOOD, DREW OUT 350ML OF BLOOD AND VENOUS PUSHED BACK. FLUID INFUSION AND METARAMINOL PUMP WERE GIVEN ALONG WITH AN EMERGENCY BLOOD TRANSFUSION, AND INTERVENTIONAL OCCLUSION TREATMENT WAS PERFORMED. IT WAS BELIEVED THAT THE PATIENT'S BLOOD VESSELS WERE SEVERELY HARDENED AND CALCIFIED, OR POSSIBLY THE STENT WAS TOO HARD AND HAD POOR FLEXIBILITY, CAUSING THE CORONARY ARTERY TO RUPTURE AND PERFORATE. NO FURTHER PATIENT INJURY REPORTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: THE PATIENT 'S VESSELS WERE THIN AND SMALL. THE PHYSICIAN ASSESSED THAT THERE WAS A POSSIBILITY THAT THE RUPTURE/PERFORATION EVENT WAS DIRECTLY RELATED TO THE USE OF THE RELEVANT DEVICE. THE PATIENT IS NOW RECOVERED AND DISCHARGED FROM HOSPITAL. TWO OF THE IMPLANTED STENTS WERE NOT IMPLANTED TO TREAT THE RUPTURE/PERFORATION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20855499 | 2024 | XIENCE SKYPOINTÃ‚Â¿ | Unknown | ['Entrapment of Device', 'Difficult to Advance', 'Device Dislodged or Dislocated', 'Material Deformation', 'No Clinical Signs, Symptoms or Conditions'] | IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT THE PROXIMAL LEFT ANTERIOR DESCENDING (LAD) CORONARY ARTERY WITH MODERATE CALCIFICATION AND MODERATE TORTUOSITY. THE 3.0X33 MM XIENCE SKYPOINT STENT DELIVERY SYSTEM (SDS) WAS BEING ADVANCED WHEN THE ENGAGEMENT OF THE GUIDING CATHETER BECAME UNSTABLE AND RESISTANCE WAS NOTED WITH THE GUIDING CATHETER DURING ADVANCEMENT. WHILE PUSHING AND PULLING THE SDS AROUND THE GUIDING CATHETER WITH RESISTANCE, THE STENT HIT THE EDGE OF THE GUIDING CATHETER CAUSING STENT DEFORMATION AND DISPLACEMENT ON THE BALLOON. THE STENT REMAINED ON THE BALLOON AND THE ENTIRE SDS WAS ULTIMATELY RETRIEVED VIA SNARE, USING A FEMORAL APPROACH. A NON-ABBOTT STENT WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT SEQUELA AND THERE WAS NO CLINICALLY SIGNIFICANT DELAY IN THE PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED. MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. A CONCLUSIVE CAUSE FOR THE REPORTED DIFFICULTIES CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. IT WAS REPORTED THAT THE STENT DELIVERY SYSTEM (SDS) MET RESISTANCE WITH A GUIDING CATHETER, RESULTING IN THE STENT BECAME DEFORMED AND DISPLACED ON THE BALLOON. FACTORS THAT MAY CONTRIBUTE TO DIFFICULTY ADVANCING THE SDS THROUGH THE GUIDING CATHETER INCLUDE, BUT ARE NOT LIMITED TO, INNER DIAMETER OF THE CATHETER, SYSTEM PROFILE, GUIDEWIRE MOVEMENT (IM COLLAPSE), INADEQUATE DELIVERABILITY, OR CHALLENGING ANATOMY. BASED ON THE INFORMATION PROVIDED AND BECAUSE THE DEVICE WAS NOT RETURNED FOR ANALYSIS, IT IS UNKNOWN WHAT MAY HAVE CAUSED THE REPORTED DIFFICULTIES DURING ADVANCEMENT. IN ADDITION, IT WAS REPORTED THAT THE SDS BECAME DEFORMED AND DISPLACED. IN THIS CASE, IT IS POSSIBLE THE MANIPULATION OF THE SDS, WHEN RESISTANCE WAS MET, CONTRIBUTED TO THE REPORTED MATERIAL DEFORMATION AND DEVICE DISLODGED OR DISLOCATED. ADDITIONALLY, IT IS LIKELY THAT THE INTERACTION WITH THE GUIDING CATHETER CONTRIBUTED TO THE ENTRAPMENT. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE.
20858289 | 2024 | SYNERGY XD | Male | ['Activation Failure', 'Chest Pain', 'Myocardial Infarction', 'Restenosis'] | AGENT IDE IT WAS REPORTED THAT RESTENOSIS OCCURRED. ON (B)(6) 2023, THE SUBJECT PRESENTED WITH MYOCARDIAL INFARCTION. A 2.5 MM X 12 MM SYNERGY DRUG ELUTING STENT (DES) WAS IMPLANTED IN THE SAPHENOUS VEIN GRAFT (SVG)-DIAGONAL GRAFT AND A 2.25 MM X 16 MM SYNERGY DES WAS IMPLANTED IN THE SVG-RAMUS GRAFT. ON (B)(6) 2024, THE SUBJECT PRESENTED TO THE EMERGENCY DEPARTMENT (ED) WITH CHEST PAIN. AT THE TIME OF EVENT THE SUBJECT WAS ON CLOPIDOGREL, WHICH WAS CONTINUED. BASED ON THE SYMPTOMS AND BIOMARKER ELEVATION THE SUBJECT WAS DIAGNOSED WITH NON-ST ELEVATED MYOCARDIAL INFARCTION (MI) AND WAS RECOMMENDED FOR CARDIOLOGY CONSULTATION. THE LOCATION OF THE MI WAS NOT IDENTIFIABLE, AND THE MI WAS CLASSIFIED AS NON-Q-WAVE MI. ON (B)(6) 2024, THE SUBJECT WAS ADMITTED TO THE HOSPITAL FOR FURTHER EVALUATION AND TREATMENT. ON (B)(6) 2024, DIAGNOSTIC CORONARY ANGIOGRAPHY REVEALED 70% STENOSIS AT THE DIAGONAL GRAFT AND 90% IN-STENT RESTENOSIS (ISR) AT THE RAMUS GRAFT. THE SUBJECT WAS RECOMMENDED FOR REVASCULARIZATION. INTRAVASCULAR ULTRASOUND (IVUS) EXAMINATION REVEALED UNDER EXPANSION OF THE PREVIOUSLY PLACED 2.25 MM X 16 MM SYNERGY DES AT THE SVG-RAMUS GRAFT. THE 90% ISR AT THE SVG-RAMUS GRAFT WAS TREATED WITH A NON-BOSTON SCIENTIFIC (NON-BSC) 2.25 MM X 15 MM STENT. POST REVASCULARIZATION, THE RESIDUAL STENOSIS WAS 0% AND TIMI FLOW WAS 3. THE SUBJECT WAS DISCHARGED WITH DAPT.
20862360 | 2024 | ONYX FRONTIER | Male | ['Device Dislodged or Dislocated', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. IT WAS REPORTED THAT ONE ONYX FRONTIER CORONARY DRUG ELUTING STENT DISLODGED IN THE VESSEL. THE CASE WAS DIFFICULT AND COMPLEX. THE STENT WAS ABLE TO BE SNARED OUT. NO FURTHER PATIENT INJURY REPORTED. ADDITIONAL INFORMATION: IT WAS NOT DIFFICULT TO REMOVE THE STYLET OR PROTECTIVE SHEATH BEFORE USE. THE 4.5X22MM AND 3.0X38MM DES USED WERE ALSO ONYX FRONTIER STENTS. THERE ARE NO COMPLAINTS AGAINST THE OTHER MEDTRONIC DEVICES USED DURING THE PROCEDURE. CORRECTION: SECTION A. PATIENT INFORMATION 5A. ETHNICITY. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: THE DISLODGEMENT OCCURRED DURING DELIVERY TO/AT LESION. THE DEVICE WAS BEING USED TO TREAT A MODERATELY TORTUOUS, MODERATELY CALCIFIED LESION WITH 100% STENOSIS IN THE PROXIMAL LEFT ANTERIOR DESCENDING (LAD) ARTERY. THE DEVICE WAS INSPECTED PRIORS TO USE WITH NO ISSUES. NEGATIVE PREP WAS NOT PERFORMED. THE LESION WAS PRE-DILATED. EXCESSIVE FORCE WAS NOT USED DURING DELIVERY. A RIGHT RADIAL APPROACH WAS USED. A 6F GUIDE WAS USED WITH A NON-MEDTRONIC GUIDEWIRE. DIRECT STENT OF THE OSTIAL RCA WAS PERFORMED FIRST WITH A 4.5X22MM DES, WITH GOOD ANGIOGRAPHIC RESULTS. THE LAD WAS THEN CROSSED AND DILATION USING A 2.5MM BALLOON. STENTING WAS SUBSEQUENTLY PERFORMED USING A 3.0X38MM DES. THE DIAGONAL WAS WIRED, AND DILATION AND STENTING WAS PERFORMED IN THE MIDPORTION. SUBSEQUENT WAS STENTING IN THE PROXIMAL PORTION WAS DONE USING ONE 2.75X18MM ONYX FRONTIER DES. THE LAD HAD 0% STENOSIS POST STENTING WITH DIFFUSE DISTAL DISEASE. THE PATIENT IS ALIVE WITH NO FURTHER INJURY. SECTION A. PATIENT INFORMATION UPDATED. SECTION B7. RELEVANT HISTORY UPDATED. INITIAL REPORTER NAME AND PHONE NUMBER PROVIDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION - IMAGE ANALYSIS: THE TARGET LESIONS IN THE LAD AND D1 WERE CLEARLY EVIDENT FROM THE IMAGES PROVIDED. NO IMAGES SHOWING THE ATTEMPTED DELIVERY AND WITHDRAWAL OF THE COMPLAINT STENT WERE SEEN IN THE IMAGES. DURING TREATMENT THE USE OF THE IVUS DEVICE CAN BE SEEN. IVUS IMAGE REVIEW ARE CAPTURED IN THE IMAGE REVIEW INVESTIGATION TASK. A LONG STENT MOST LIKELY THE 3.0X38MM ONYX FRONTIER STENT WAS SUCCESSFULLY COMPLETED. THE DEPLOYMENT OF THE 2.75X18MM ONYX FRONTIER STENT WAS COMPLETED IN THE D1. THE SUCCESSFUL STENT DEPLOYMENT IN THE RCA CAN ALSO BE CONFIRMED. SINCE NO IMAGES WERE PROVIDED OF THE DISLODGEMENT AND RETRIEVAL OF THE ONYX FRONTIER STENT THE CAUSE OF THE ISSUE CANNOT BE CONFIRMED FROM THE IMAGES PROVIDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL IMAGE ANALYSIS: IVUS IMAGES: UNABLE TO DISTINGUISH WHICH VESSELS WERE IMAGED. MEASUREMENTS WERE NOTED, ALONG WITH MODERATE CALCIFICATION. NO EVIDENCE OF DISSECTIONS. ANGIOGRAPHIC IMAGES: MULTIVESSEL DIFFUSE CORONARY ARTERY DISEASE (CAD) WITH TORTUOSITY FOR THE LEFT ANTERIOR DESCENDING (LAD) AND BRANCHES. INTERVENTION OF DIRECT STENTING FOR THE RIGHT CORONARY ARTERY (RCA) WERE NOTED WITH NO COMPLICATIONS. THE LAD WAS WIRED WITH A SHORT BALLOON INTRODUCED INTO THE PROXIMAL LAD. STENT DELIVERED INTO THE PROXIMAL/MID LAD AND ANOTHER GUIDEWIRE IN THE DIAGONAL 1 (D1). BALLOON DELIVERED TO THE MID PORTION OF D1. POST DILATATION OF THE LAD PERFORMED. NO ANGIOGRAPHIC COMPLICATIONS. NEXT IMAGES PREVIEW WHAT APPEARS TO BE A SHORT BALLOON AT THE OSTIUM OF THE LAD. A GUIDE EXTENSION CATHETER WAS USED AT THE OSTIUM OF THE LAD. SHORT BALLOON INFLATION AT THE OSTIAL LAD. STENT DELIVERED TO THE D1, AND A BALLOON DELIVERED WITHIN THE LAD STENT, THEN BOTH INFLATED. NO COMPLICATIONS NOTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20862388 | 2024 | ONYX FRONTIER | Male | ['Difficult to Advance', 'Device Dislodged or Dislocated', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. IT WAS REPORTED THAT ONE ONYX FRONTIER CORONARY DRUG ELUTING STENT DISLODGED IN THE VESSEL. THE CASE WAS DIFFICULT AND COMPLEX. THE STENT WAS SUCCESSFULLY SNARED OUT. NO FURTHER PATIENT INJURY REPORTED. ADDITIONAL INFORMATION: THE PATIENT PRESENTED TO HOSPITAL WITH A NON-ST-ELEVATION MYOCARDIAL INFARCTION (NSTEMI). THE DISLODGEMENT OCCURRED DURING DELIVERY TO/AT LESION. THE STENT WAS BEING USED TO TREAT A MODERATELY TORTUOUS, MODERATELY CALCIFIED LESION WITH 90% STENOSIS IN THE PROXIMAL DIAGONAL BRANCH (D1). THERE WAS TIMI 3 FLOW. THE DEVICE WAS INSPECTED PRIOR TO USE WITH NO ISSUES. NEGATIVE PREP WAS NOT PERFORMED. THE LESION WAS PRE-DILATED. RESISTANCE WAS ENCOUNTERED WHEN ADVANCING THE DEVICE. EXCESSIVE FORCE WAS NOT USED. A 6F GUIDE CATHETER CANNULATED THE LEFT MAIN (LM) AND TWO NON-MEDTRONIC (MDT) GUIDEWIRES CROSSED THE D1 LESION. A 2.5MM COMPLIANT BALLOON PRE-DILATED THE PROXIMAL D1 STENOSIS TO 8 ATM. INTRAVASCULAR ULTRASOUND (IVUS) WAS USED WHICH DEMONSTRATED DISEASE ALL THE WAY TO THE OSTIUM. A 2.0X26MM ONYX FRONTIER STENT WAS DEPLOYED FROM THE PROXIMAL TO MID D1 AT 12 ATM. FOLLOWING THIS AN ATTEMPT WAS MADE TO ADVANCE THE 3.0X8MM ONYX FRONTIER STENT TO THE OSTIUM OF THE D1, OVERLAPPING THE PROXIMAL EDGE OF THE INITIAL STENT. HOWEVER, THERE WAS SIGNIFICANT DIFFICULTY ADVANCING THE STENT BEYOND THE OSTIAL PORTION OF THE D1, AND DURING THIS PROCESS THE STENT STRIPPED OFF THE BALLOON WITHOUT DEPLOYMENT. AN ATTEMPT WAS MADE TO USE THE BALLOON TO RE-CAPTURE THE DISLODGED STENT INTO THE GUIDE, BUT THIS WAS NOT SUCCESSFUL. A LEFT ANTERIOR DESCENDING (LAD) ARTERY WIRE WAS ATTEMPTED TO WRAP AROUND THE STENT TO PULL IT OUT, BUT THIS WAS ALSO UNSUCCESSFUL. GUIDE AND WIRE ACCESS WERE THEN LOST, AND THE STENT WAS FREE FLOATING IN THE PROXIMAL LAD, LEADING INTO THE OSTIAL D1. A NEW NON-MDT GUIDEWIRE WAS USED AND A 4-8MM NON-MDT SNARE WAS USED TO SUCCESSFULLY SNARE THE DISLODGED STENT. THEN, ANOTHER 3.0X8MM ONYX FRONTIER STENT WAS ADVANCED TO THE D1, HOWEVER, THERE WAS DIFFICULTY CROSSING INTO THE PROXIMAL D1. THERE WAS CONCERN OF A POSSIBLE CORONARY DISSECTION AT THE OSTIUM OF THE D1 DUE TO THE PREVIOUSLY DEPLOYED STENT. EVENTUALLY, IT WAS POSSIBLE TO ADVANCE A 2.5MM NC BALLOON TO THE MID D1 FOR POST DILATION OF THE 2.0X26MM ONYX FRONTIER STENT, AS WELL AS PRE-DILATION OF THE OSTIUM D1 TO A MAXIMUM OF 18 ATM. THE SECOND 3.0X8MM ONYX FRONTIER STENT WAS THEN SUCCESSFULLY ADVANCED TO THE OSTIAL D1 OVERLAPPING THE PROXIMAL EDGE OF THE 2.0X26MM ONYX FRONTIER STENT AND WAS SUCCESSFULLY DEPLOYED TO 14 ATM. IVUS AFTER STENT DEPLOYMENT DEMONSTRATED A WELL-APPOSED STENT, HOWEVER, THERE WAS MILD UNDER-EXPANSION OF THE 3.0X8MM ONYX FRONTIER STENT. A 3.0MM NC BALLOON WAS THEN USED TO POST DILATE THE 3.0X8MM ONYX FRONTIER STENT TO 20 ATM. IVUS SHOWED GOOD STENT EXPANSION AND APPOSITION, WITH NO EVIDENCE OF STENT EDGE DISSECTION. POST INTERVENTION THERE WAS TIMI FLOW 3, AND 0% RESIDUAL STENOSIS. THE PATIENT IS ALIVE WITH NO FURTHER INJ URY. SECTION A. PATIENT INFORMATION UPDATED. SECTION B7. RELEVANT HISTORY UPDATED. INITIAL REPORTER NAME AND PHONE NUMBER PROVIDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: IT WAS NOT DIFFICULT TO REMOVE THE STYLET OR PROTECTIVE SHEATH BEFORE USE. THERE ARE NO COMPLAINTS AGAINST THE OTHER MEDTRONIC DEVICES USED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. IMAGE ANALYSIS: PROCEDURAL IMAGES WERE PROVIDED FOR REVIEW CONFIRMING THE TARGET LESION IN THE D1. MODERATE TORTUOSITY AND MILD CALCIFICATION WERE PRESENT IN THE LAD AND DIAGONAL BRANCH. THERE WAS STENOSIS AT THE OSTIUM OF THE DIAGONAL AND IN THE MID-SECTION OF THE BRANCH. THE VESSEL WAS PRE-DILATED WITHOUT EVENT. NO ANGIOGRAPHIC COMPLICATIONS NOTED. THIS WAS FOLLOWED BY SUCCESSFUL DEPLOYMENT OF A STENT IN THE PROXIMAL TO MID D1. IVUS CATHETER WAS UTILIZED BUT NO IMAGES WERE PROVIDED. SUBSEQUENT POST DILATIONS WITH A MUCH SHORTER BALLOON WERE DONE, WITH NO ISSUES NOTED. THE OSTIAL VESSEL WAS NOT TREATED AND REMAINED STENOTIC. THE FOLLOWING IMAGES SHOW A DISLODGED STENT, BUT NO IMAGES WERE PROVIDED SHOWING THE ATTEMPTED DELIVERY OF THE STENT. THE DISLODGMENT SEQUENCE IS NOT VISIBLE ON THE PROVIDED ANGIOGRAMS, BUT THE DISLODGED STENT, INTENDED FOR THE PROXIMAL DIAGONAL, IS VISUALIZED AS WELL AS THE SUCCESSFUL SNARING AND REMOVAL, WITH NO ANGIOGRAPHIC COMPLICATIONS. THE STENT WAS REMOVED FROM THE VESSEL FOLLOWING POST DILATION OF THE NEWLY DEPLOYED STENT. ANOTHER SHORT OSTIAL STENT WAS DELIVERED AND DEPLOYED RESULTING IN A FULLY PATENT VESSEL. THE MOST LIKELY CAUSE OF THE DISLODGEMENT WAS DUE TO THE VESSEL STENOSIS BUT THIS CANNOT BE CONFIRMED AS THE MECHANISM OF DISLODGEMENT WAS NOT PROVIDED ON THE IMAGES. THE REPORTED POSITIONING DIFFICULTIES WERE NOT SEEN ON THE IMAGES PROVIDED, HOWEVER, THE POSSIBLY EXISTS BUT WAS NOT CONFIRMED THAT THE LESION MORPHOLOGY IMPACTED ON THE FAILURE TO DELIVER. THERE WAS NO EVIDENCE OF STENT MALAPPOSITION AS THE ON THE IMAGES PROVIDED. TIMI3 FLOW BEFORE DURING AND AFTER INTERVENTIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20865513 | 2024 | SYNERGY XD | Unknown | ['Break', 'Positioning Failure', 'Difficult to Remove', 'Material Rupture', 'Difficult to Advance', 'Cardiac Arrest', 'Reduced Blood Flow'] | IT WAS REPORTED THAT A BALLOON ISSUE OCCURRED RESULTING IN PARTIAL STENT DEPLOYMENT. THE 80% STENOSED TARGET LESION WAS LOCATED IN THE MODERATELY TORTUOUS AND SEVERELY CALCIFIED MID-DISTAL LEFT ANTERIOR DESCENDING (LAD) ARTERY. ROTATIONAL ATHERECTOMY WAS PERFORMED AND A SCORING BALOON AND 2.00 NC BALLOON WERE USED TO PREPARE THE LESION., A 2.50MM X 48MM SYNERGY XD DRUG-ELUTING STENT WAS ADVANCED BUT COULD NOT BE DELIVERED TO THE LESION AREA. A 2.00 X 15 MM NC EMERGE BALLOON CATHETER WAS ADVANCED TO FURTHER PRE-DILATE THE LESION, AND AN IVUS RUN WAS CONDUCTED. THE 2.50MM X 48MM SYNERGY XD WAS RE-INTRODUCED, BUT DURING DEPLOYMENT AT 10 ATMOSPHERES; THE BALLOON RUPTURED. THE PROXIMAL PORTION OF THE STENT WAS PARTIALLY DEPLOYED, WHILE THE DISTAL PART REMAINED INTACT. ATTEMPTS TO REMOVE THE STENT USING A GUIDE EXTENSION CATHETER WERE UNSUCCESSFUL, NECESSITATING MULTIPLE SHOCKS TO THE PATIENT DUE TO NO FLOW. THE PATIENT WENT INTO CARDIAC ARREST AND WAS SUBSEQUENTLY PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). WHEN THE OPERATOR TRIED TO REMOVE THE STENT MANUALLY, THE SHAFT BROKE. A GUIDE EXTENSION CATHETER WAS INTRODUCED AND A 2.5MM NC BALLOON WAS USED TO ANCHOR THE RUPTURED BALLOON. THE WHOLE SYSTEM WAS THEN REMOVED TOGETHER. THE PROCEDURE PROCEEDED WITH LAD STENTING AND AN INTRA-AORTIC BALLOON PUMP (IABP) WAS INSERTED IN PLACE OF THE ECMO. TWO DAYS LATER THE PATIENT UNDERWENT SURGERY TO STOP BLEEDING AT THE BRACHIAL SITE AND IS CURRENTLY ON DIALYSIS DUE TO KIDNEY FAILURE. IT WAS REPORTED THAT A BALLOON ISSUE OCCURRED RESULTING IN PARTIAL STENT DEPLOYMENT. THE 80% STENOSED TARGET LESION WAS LOCATED IN THE MODERATELY TORTUOUS AND SEVERELY CALCIFIED MID-DISTAL LEFT ANTERIOR DESCENDING (LAD) ARTERY. ROTATIONAL ATHERECTOMY WAS PERFORMED AND A SCORING BALOON AND 2.00 NC BALLOON WERE USED TO PREPARE THE LESION., A 2.50MM X 48MM SYNERGY XD DRUG-ELUTING STENT WAS ADVANCED BUT COULD NOT BE DELIVERED TO THE LESION AREA. A 2.00 X 15 MM NC EMERGE BALLOON CATHETER WAS ADVANCED TO FURTHER PRE-DILATE THE LESION, AND AN IVUS RUN WAS CONDUCTED. THE 2.50MM X 48MM SYNERGY XD WAS RE-INTRODUCED, BUT DURING DEPLOYMENT AT 10 ATMOSPHERES; THE BALLOON RUPTURED. THE PROXIMAL PORTION OF THE STENT WAS PARTIALLY DEPLOYED, WHILE THE DISTAL PART REMAINED INTACT. ATTEMPTS TO REMOVE THE STENT USING A GUIDE EXTENSION CATHETER WERE UNSUCCESSFUL, NECESSITATING MULTIPLE SHOCKS TO THE PATIENT DUE TO NO FLOW. THE PATIENT WENT INTO CARDIAC ARREST AND WAS SUBSEQUENTLY PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). WHEN THE OPERATOR TRIED TO REMOVE THE STENT MANUALLY, THE SHAFT BROKE. A GUIDE EXTENSION CATHETER WAS INTRODUCED AND A 2.5MM NC BALLOON WAS USED TO ANCHOR THE RUPTURED BALLOON. THE WHOLE SYSTEM WAS THEN REMOVED TOGETHER. THE PROCEDURE PROCEEDED WITH LAD STENTING AND AN INTRA-AORTIC BALLOON PUMP (IABP) WAS INSERTED IN PLACE OF THE ECMO. TWO DAYS LATER THE PATIENT UNDERWENT SURGERY TO STOP BLEEDING AT THE BRACHIAL SITE AND IS CURRENTLY ON DIALYSIS DUE TO KIDNEY FAILURE. IT WAS FURTHER REPORTED THAT ALSO WITH THE SAME PATIENT, A 90% STENOSED TARGET LESION LOCATED IN THE SEVERELY TORTUOUS AND SEVERELY CALCIFIED LEFT ANTERIOR DESCENDING ARTERY WAS TREATED. DURING THE PROCEDURE, THREE NC EMERGE BALLOON CATHETERS (1.50 MM X 15 MM, 2.50 MM X 15 MM, AND 3.50 MM X 15MM) WERE UNABLE TO CROSS THE TARGET LESION, AND THREE NC EMERGE BALLOON CATHETER BALLOONS (2.00MM X 15MM, 2.00MM X 15MM, AND 2.25 MM X 8 MM) RUPTURED, ALL DUE TO HEAVY CALCIFICATION. NO DEVICE FRAGMENTS WERE LEFT IN THE PATIENT. THE SYNERGY XD 2.50/48MM WAS RETURNED FOR ANALYSIS IN TWO SECTIONS DUE TO A BREAK AT THE GUIDEWIRE EXIT PORT. A VISUAL AND TACTILE EXAMINATION IDENTIFIED NO KINKS OR DAMAGES ALONG THE ENTIRE LENGTH OF THE HYPOTUBE. A VISUAL AND TACTILE EXAMINATION OF THE DISTAL EXTRUSION IDENTIFIED A BREAK IN THE MIDSHAFT EXTRUSION AT THE SITE OF THE GUIDEWIRE EXIT PORT. A VISUAL EXAMINATION OF THE BALLOON IDENTIFIED NO TEARS, ALTHOUGH THERE WERE TRACES OF BLOOD PRESENT INSIDE THE BALLOON. A VISUAL AND TACTILE EXAMINATION OF THE STENT IDENTIFIED THAT THE STENT WAS PULLED DISTALLY ON THE BALLOON AND WAS STRETCHED BEYOND THE DISTAL TIP OF THE DEVICE. THE PROXIMAL EDGE OF THE STENT WAS BUNCHED AND POSITIONED TOWARDS THE DISTAL MARKER BAND. A MICROSCOPIC EXAMINATION OF THE BREAK SITE AT THE GUIDEWIRE EXIT PORT NOTED THAT THE EXTRUSION WAS STRETCHED. A MICROSCOPIC EXAMINATION OF THE BALLOON MATERIAL IDENTIFIED NO TEARS OR PINHOLES. A MICROSCOPIC EXAMINATION OF THE CRIMPED STENT IDENTIFIED THAT THE STENT WAS STRETCHED DISTALLY, WITH STENT STRUTS STRETCHED, BUNCHED, AND MISALIGNED. THE STENT WAS ATTACHED TO THE BALLOON JUST PROXIMAL TO THE DISTAL MARKER BAND, WITH THE MAIN BODY OF THE STENT PULLED BEYOND THE TIP SECTION OF THE DEVICE. A MICROSCOPIC EXAMINATION OF THE PROXIMAL AND DISTAL MARKER BANDS IDENTIFIED NO DAMAGE. DUE TO THE STRETCHING IN THE INNER LUMEN, THE DISTANCE FROM THE PROXIMAL MARKER BAND TO THE DISTAL MARKER BAND MEASURED 52MM (DEVICE SPECIFICATION IS 48MM). A MICROSCOPIC EXAMINATION OF THE TIP IDENTIFIED NO DAMAGES. DUE TO THE BUNCHING AT THE DISTAL END OF THE 6 FRENCH INTRODUCER SHEATH, IT WAS NOT POSSIBLE TO RETRACT THE BROKEN DISTAL SECTION OF THE SYNERGY XD DEVICE BACK THROUGH THE INTRODUCER. INSTEAD, THE BROKEN SECTION OF THE SYNERGY XD DEVICE WAS PULLED DISTALLY FROM THE INTRODUCER TO REMOVE IT. IT WAS POSSIBLE TO REMOVE THE TWO GUIDEWIRES AND THE NC 2.5 X 20MM BALLOON CATHETER FROM THE GUIDE WITH NO RESISTANCE. DUE TO THE BREAK IN THE EXTRUSION, IT WAS NOT POSSIBLE TO INFLATE THE BALLOON USING NORMAL METHODS. HOWEVER, BY DISSECTING THE DISTAL EXTRUSION APPROXIMATELY 6 CM FROM THE PROXIMAL BALLOON SLEEVE AND USING A FLUSHING NEEDLE, IT WAS POSSIBLE TO INJECT SOME LIQUID INTO THE BALLOON. ALTHOUGH NO LEAK WAS EVIDENT IN THE BALLOON, THE PRESSURE COULD NOT REACH THE NOMINAL PRESSURE OF 11 ATMOSPHERES. MEDIA PROVIDED BY THE CUSTOMER WAS REVIEWED BY BOSTON SCIENTIFIC MEDICAL PERSONNEL. THE CLINICAL DESCRIPTION ALIGNED WITH THE ANGIOGRAPHIC IMAGES. THE PATIENT HAD COMPLEX CORONARY DISEASE. THE LAD HAS A LONG SEGMENT OF SEVERE STENOSIS. SEVERE ANGIOGRAPHIC CALCIFICATION IS NOTED. THERE ARE ALSO CHRONIC TOTALLY OCCLUDED (CTO) LESIONS IN THE LEFT CIRCUMFLEX (LCX) AND RIGHT CORONARY ARTERY (RCA). THE INTERVENTION SHOWED EITHER BALLOON UNCROSSABLE AND/OR UNDILATABLE DISEASE IN THE MORE DISTAL SEGMENT. A MID MARKER BALLOON AND GUIDE CATHETER EXTENSION WERE SEEN. ROTATIONAL ATHERECTOMY WAS CARRIED OUT. REDUCED FLOW IS APPARENT AFTER THE ATHERECTOMY DEVICE HAD CROSSED THE SEGMENT. MULTIPLE PRE-DILATIONS WERE CARRIED OUT WITH SMALL DIAMETER BALLOONS AND FLOW WAS RESTORED. A STENT WAS POSITIONED AND AFTER DEPLOYMENT STARTED, A BALLOON RUPTURE WAS EVIDENT EARLY WITH CONTRAST EXTRUDING INTO THE SEPTAL BRANCH AT THE PROXIMAL EDGE OF THE STENT. ATTEMPTS TO WITHDRAW THE BALLOON AND STENT WERE CAPTURED. THERE WAS MARKED SHORTENING OF THE STENT. THERE WAS PROBABLY A PROXIMAL DEFORMATION WHEN THE GUIDE CATHETER EXTENSION (GCE) WAS USED TO TRY TO CAPTURE THE STENT AND BALLOON. AFTER PERIPHERAL ACCESS WAS OBTAINED, THE IMAGES SHOW A DEFORMED AND SHORTENED STENT AND ALSO NO FLOW TO THE APICAL LAD. THE REVIEW THEN NOTED THAT THE STENT HAD BEEN WITHDRAWN TO THE RADIAL ARTERY. THE LAD IS ULTIMATELY RE-WIRED AND STENTED, WITH POST DILATION. SEPTAL BRANCHES ARE ALSO RESCUED. THE END RESULT IS TIMI3 FLOW WITH A STENTED LAD SEGMENT. THE REVIEW CONCLUDED THAT UNDERLYING CALCIFICATION WAS A FACTOR IN THE BALLOON RUPTURE. IT WAS ANTICIPATED THAT THIS LED TO INTERACTION BETWEEN THE BALLOON MATERIAL AND STENT ITSELF THAT PREVENTED THE BALLOON FROM BEING REMOVED FROM WITHIN THE STENT AND THIS LED TO THE SUBSEQUENT SPIRAL OF EVENTS.
20866780 | 2024 | XIENCE SKYPOINTÃ‚Â¿ | Unknown | ['Entrapment of Device', 'Material Rupture', 'Activation Failure', 'No Clinical Signs, Symptoms or Conditions'] | MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B3: DATE OF EVENT IS ESTIMATED AS 05/01/2024. D4: A PARTIAL UDI WAS PROVIDED AS THE LOT NUMBER IS NOT KNOWN. MEDSUN REPORT (B)(4) RECEIVED WHICH STATES: DESCRIBE THE EVENT OR PROBLEM: ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION (STEMI) PATIENT FROM EMERGENCY ROOM (ER), PERFORMING PERCUTANEOUS CORONARY INTERVENTION (PCI), WENT IN WITH XIENCE SKYPOINT DRUG ELUTING STENT. WENT TO DEPLOY THE STENT. THE BALLOON WOULD NOT KEEP ITS PRESSURE UP. THE BALLOON HAD RUPTURED WHILE GOING UP. DOCTOR EXPLAINED THAT THE BALLOON BURSTS DURING INFLATION IN THE RIGHT CORONARY ARTERY (RCA) AND HE WAS HAVING DIFFICULTY REMOVING THE DEVICE VIA RIGHT (RT) RADIAL ACCESS. WE DECIDED TO GET FEMORAL ACCESS AND PROCEED TO REMOVE THE DEVICE WITH ENSNARE. THE DEVICE WAS EVENTUALLY REMOVED. THE VENDOR WAS CONTACTED, AND THEY ARE REQUESTING TO PICK UP THE DEVICE FOR INSPECTION. A REVIEW OF THE CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED TWO RELATED COMPLAINT ASSESSMENT CAPAS ISSUE RELATED TO MATERIAL RUPTURE. IN THIS CASE, THE REPORTED DEVICE WAS NOT RETURNED; THEREFORE, A SIMILARITY BETWEEN THE REPORTED ISSUE AND THE COMPLAINT ASSESSMENTS CANNOT BE DETERMINED. BASED ON THESE REVIEWS, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE LOT NUMBER WAS NOT REPORTED. THE INVESTIGATION WAS UNABLE TO DETERMINE A CONCLUSIVE CAUSE FOR THE REPORTED DIFFICULTIES; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO OPERATIONAL CIRCUMSTANCES. IT WAS REPORTED THE STENT DELIVERY SYSTEM (SDS) BALLOON RUPTURED DURING INFLATION AND FAILED TO ACTIVATE. FACTORS THAT MAY CONTRIBUTE TO THE REPORTED BALLOON MATERIAL RUPTURE INCLUDE, BUT ARE NOT LIMITED TO, MATERIAL DAMAGE, INTERACTIONS WITH OTHER DEVICES OR INTERACTION WITH LESION CALCIFICATION AND TORTUOSITY. BASED ON THE INFORMATION PROVIDED AND BECAUSE THE DEVICE WAS NOT RETURNED FOR ANALYSIS, IT IS UNKNOWN WHAT MAY HAVE CAUSED THE REPORTED MATERIAL RUPTURE. ADDITIONALLY, IT IS POSSIBLE THAT THE REPORTED SDS MATERIAL RUPTURE CONTRIBUTED TO THE REPORTED ACTIVATION FAILURE. A STENT THAT FAILED TO ACTIVATE MAY CAUSE RESISTANCE DURING REMOVAL, POSSIBLY CONTRIBUTING TO THE REPORTED ENTRAPMENT. THE ENTRAPPED SDS WAS RETRIEVED VIA A SNARE. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. E1: INITIAL REPORTER UPDATED FROM RISK MANAGER: (B)(6) TO UNKNOWN PHYSICIAN.
20867152 | 2024 | XIENCE SKYPOINTÃ‚Â¿ | Unknown | ['Improper or Incorrect Procedure or Method', 'Failure to Advance', 'Device Dislodged or Dislocated', 'Device Embedded In Tissue or Plaque'] | THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED DIFFICULTIES AND PATIENT EFFECTS, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT THE PROXIMAL CIRCUMFLEX CORONARY ARTERY WITH MODERATE CALCIFICATION, MILD TORTUOSITY AND 90% STENOSIS. THE 3.5X38 MM XIENCE SKYPOINT STENT DELIVERY SYSTEM (SDS) COULD NOT ADVANCE TO THE INTENDED LOCATION DUE TO PREVIOUSLY IMPLANTED STENTS; THEREFORE, THE SDS WAS REMOVED AND A 3.5X33 MM XIENCE SDS WAS IMPLANTED. NEXT, THE ARTERY WAS INTENDED TO BE TREATED WITH THE SAME 3.5X38 MM XIENCE SKYPOINT SDS AND WAS RE-INSERTED. DURING ADVANCEMENT THERE WAS RESISTANCE NOTED RESULTING IN THE FAILURE TO ADVANCE AND THE STENT DISLODGING IN THE LESION. FACTORS THAT MAY CONTRIBUTE TO THE RESISTANCE ENCOUNTERED INCLUDE, BUT ARE NOT LIMITED TO, OUTER DIAMETER OF THE STENT DELIVERY SYSTEM, DEVICE SUPPORT, BUILDUP OF PROCEDURAL CONTAMINANTS, USER TECHNIQUE, AND/OR DAMAGE TO THE STENT DELIVERY SYSTEM. BASED ON THE INFORMATION PROVIDED AND BECAUSE THE DEVICE WAS NOT RETURNED FOR ANALYSIS, IT IS UNKNOWN WHAT MAY HAVE CAUSED THE REPORTED RESISTANCE ENCOUNTERED. ADDITIONALLY, IT IS POSSIBLE THAT MANIPULATION OF THE DEVICE, WHEN RESISTANCE WAS MET, CONTRIBUTED TO THE REPORTED DISLODGEMENT. IT WAS REPORTED THAT THE DISLODGED STENT WAS EMBEDDED IN THE PATIENT'S ANATOMY. IT SHOULD BE NOTED THAT THE XIENCE SKYPOINT INSTRUCTIONS FOR USE STATES: AN UNEXPANDED STENT MAY BE RETRACTED INTO THE GUIDING CATHETER ONE TIME ONLY. AN UNEXPANDED STENT SHOULD NOT BE REINTRODUCED INTO THE ARTERY ONCE IT HAS BEEN PULLED BACK INTO THE GUIDING CATHETER. SUBSEQUENT MOVEMENT IN AND OUT THROUGH THE DISTAL END OF THE GUIDING CATHETER SHOULD NOT BE PERFORMED, AS THE STENT MAY BE DAMAGED OR DISLODGED DURING RETRACTION BACK INTO THE GUIDING CATHETER. IN THIS CASE IT IS UNKNOWN IF THE IFU DEVIATION DIRECTLY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. D4: A PARTIAL UDI WAS PROVIDED AS THE LOT NUMBER IS NOT KNOWN. H6: MEDICAL DEVICE PROBLEM CODE 2017: RE-INSERTION. IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT THE PROXIMAL CIRCUMFLEX CORONARY ARTERY WITH MODERATE CALCIFICATION, MILD TORTUOSITY AND 90% STENOSIS. THE 3.5X38 MM XIENCE SKYPOINT STENT DELIVERY SYSTEM (SDS) COULD NOT ADVANCE TO THE INTENDED LOCATION DUE TO PREVIOUSLY IMPLANTED STENTS; THEREFORE THE SDS WAS REMOVED AND A 3.5X33 MM XIENCE SDS WAS IMPLANTED. NEXT, THE PROXIMAL CIRCUMFLEX CORONARY ARTERY WAS INTENDED TO BE TREATED AND THE SAME 3.5X38 MM XIENCE SKYPOINT SDS WAS RE-ADVANCED. DURING ADVANCEMENT THERE WAS RESISTANCE WHEN THE DEVICE REACHED THE LESION AND THE STENT DISLODGED FROM THE DELIVERY SYSTEM AND ATTEMPTS TO SNARE OR TRAP THE STENT WERE UNSUCCESSFUL. THE PROCEDURE WAS ABORTED AND THE PATIENT WAS HOSPITALIZED UNTIL (B)(6) 2024 WHEN AN ADDITIONAL PROCEDURE WAS PERFORMED FOR TREATMENT. THE STENT WAS ULTIMATELY CRUSHED WITH ANOTHER STENT WHERE IT IS LOCATED. THE PATIENT FINAL OUTCOME IS GOOD AND THE PATIENT IS EXPECTED TO BE DISCHARGED. NO ADDITIONAL INFORMATION WAS PROVIDED.
20877259 | 2024 | ONYX FRONTIER | Female | ['Device Damaged by Another Device', 'Device Dislodged or Dislocated', 'Positioning Problem', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. IT WAS REPORTED THAT DURING A PERCUTANEOUS CORONARY INTERVENTION (PCI) PROCEDURE ONE ONYX FRONTIER CORONARY DRUG ELUTING STENT DISLODGED DURING REMOVAL FOLLOWING A FAILED DELIVERY. THE LESION BEING TREATED WAS SEVERELY CALCIFICATION, MODERATELY TORTUOUS WITH 80-90% STENOSIS IN THE DISTAL OBTUSE MARGINAL (OM). THE LEFT ANTERIOR DESCENDING ARTERY (LAD) HAD SEVERE CALCIFICATION, DIAGONAL 1 HAD MID 80-90% STENOSIS, OTHERWISE MILD DISEASE THROUGHOUT. THERE WAS SEVERE CALCIFICATION IN THE LEFT MAIN (LM) WITH A PREVIOUS STENT THAT HAD IN-STENT RESTENOSIS (ISR) IN THE DISTAL SEGMENT. THERE WAS ALSO SEVERE CALCIFICATION IN THE LEFT CIRCUMFLEX (LCX) ARTERY WITH 90% OSTIAL LCX ISR, MID LCX 60% STENOSIS, OM2 WITH 80-90% STENOSIS, OTHERWISE MILD DISEASE THROUGHOUT. THE STENT WAS ATTEMPTED TO BE DELIVERED TO THE DISTAL REGION BUT COULD NOT BE DELIVERED. UPON REMOVAL OF THE DEVICE, IT WAS NOTED THAT THE STENT HAD BEEN STRIPPED OFF AND WAS INSIDE THE GUIDE. A GOOSENECK SNARE WAS USED TO SUCCESSFULLY REMOVE THE DISLODGED STENT FROM THE LM INTACT. IT WAS NO TED THAT THE PREVIOUSLY PLACED OSTIAL LM STENT HAD STRUTS HANGING OUT. A 4.5X8MM NON-MEDTRONIC BALLOON AND A 4.5X8 NON-MEDTRONIC NC BALLOON WERE USED FOR PROXIMAL OPTIMIZATION TECHNIQUE (POT) OF THE PREVIOUSLY IMPLANTED STENT, WHICH IS NOW BETTER EXPANDED, WELL APPOSED AND WELL SIZED. THE DEVICE WAS INSPECTED PRIOR TO USE WITH NO ISSUES. NEGATIVE PREP WAS NOT PERFORMED. THE LESION WAS PRE-DILATED WITH A 2.5X12MM NC BALLOON. THE DEVICE DID PASS THROUGH A PREVIOUSLY DEPLOYED STENT. RESISTANCE WAS NOT ENCOUNTERING WHEN ADVANCING THE DEVICE. EXCESSIVE FORCE WAS NOT USED. FINAL ANGIOGRAM DEMONSTRATED EXCELLENT ANGIOGRAPHY RESULT WITH TIMI FLOW III. THERE WAS 0% RESIDUAL STENOSIS. THERE WERE NO EDGE DISSECTIONS. THE PATIENT IS ALIVE WITH NO FURTHER INJURY. ADDITIONAL INFORMATION: RESISTANCE WAS NOT NOTED DURING WITHDRAWAL OF THE DEVICE. THE PREVIOUSLY PLACED OSTIAL LM STENT WITH STRUTS HANGING OUT WAS A NON-MEDTRONIC STENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20877443 | 2024 | ONYX FRONTIER | Male | ['Device Dislodged or Dislocated', 'Positioning Problem', 'Device Embedded In Tissue or Plaque', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. IT WAS REPORTED THAT DURING A PROCEDURE ONE 2.5X26MM ONYX FRONTIER CORONARY DRUG ELUTING STENT DISLODGED DURING REMOVAL FOLLOWING A FAILED DELIVERY. THE STENT WAS BEING USED TO TREAT A MODERATELY TORTUOUS, SEVERELY CALCIFIED LESION WITH 90% STENOSIS IN THE DISTAL SUPERIOR OBTUSE MARGINAL (OM). A 7F GUIDE CATHETER WAS ENGAGED IN THE LEFT MAIN (LM) AND MULTIPLE NON-MEDTRONIC GUIDEWIRES WERE ADVANCED. THE SUPERIOR OM2 BRANCH WAS PRE-DILATED SERIALLY WITH A 2.0X20MM EUPHORA BALLOON. ONE 2.5X18MM ONYX FRONTIER STENT WAS SUCCESSFULLY DEPLOYED ACROSS THE LESION. THEN THE 2.5X26MM ONYX FRONTIER WAS ATTEMPTED TO BE ADVANCED BUT COULD NOT BE ADVANCED FULLY TO THE TARGET LESION, THEREFORE THE DEVICE WAS BEING REMOVED. THE DEVICE DID PASS THROUGH A PREVIOUSLY DEPLOYED STENT. RESISTANCE WAS NOT ENCOUNTERED WHEN ADVANCING THE DEVICE. EXCESSIVE FORCE WAS NOT USED. UPON PULLING THE STENT BACK TO THE GUIDE CATHETER THE STENT DISLODGED OFF THE BALLOON INTO THE LM EXTENDING INTO THE OSTIAL CIRCUMFLEX (CX) AREA. RESISTANCE WAS NOT FELT WHEN PULLING BACK THE STENT. THE DISLODGED STENT WAS NOT REMOVED. IT WAS DECIDED THAT IT WAS SAFEST TO DEPLOY THE STENT AT THE LOCATION OF THE DISLODGEMENT GIVEN THE AMOUNT OF CALCIUM AND A PREVIOUS STENT THAT WAS LOCATED IN THE LM. THE DISLODGED STENT WAS SERIALLY DILATED WITH A 1.5X15MM NON-MEDTRONIC BALLOON, FOLLOWED BY A 2.0X15MM, 2.5X15MM, 3.5X15MM AND A EUPHORA 3.5X12MM BALLOON. THE PROXIMAL EDGE OF THE STENT WAS FURTHER DILATED WITH A 4.0X8MM BALLOON. THEN A 2.5X22MM ONYX FRONTIER STENT WAS ADVANCED PAST THE SUPERIOR OM2 STENT AND PLACED DISTALLY, OVERLAPPING THE 2.5X26MM ONYX FRONTIER STENT. INTRAVASCULAR ULTRASOUND (IVUS) WAS USED FOR VESSEL AND STENT ASSESSMENT. THE OVERLAP BETWEEN THE TWO STENTS WAS FURTHER POST DILATED WITH A EUPHORA 3.5X12MM BALLOON FOLLOWED BY A 4.5X8MM NC EUPHORA BALLOON IN THE PROXIMAL EDGE OF THE LM-LCX STENT. FINAL ANGIOGRAM DEMONSTRATED EXCELLENT ANGIOGRAPHIC RESULT WITH TIMI3 FLOW, NO PERFORATION, NO DISSECTION AND NO RESIDUAL STENOSIS. THE PATIENT IS ALIVE WITH NO FURTHER INJURY. ADDITIONAL INFORMATION: THERE IS NO COMPLAINT AGAINST ANY OF THE OTHER MEDTRONIC DEVICES USED DURING THE PROCEDURE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. PROCEDURAL IMAGE ANALYSIS: THE TARGET LESION WAS CONFIRMED TO BE PRESENT IN THE OM OF THE LCX. THE VESSEL WAS PRE-DILATED AND WHAT APPEARED TO BE THE 2.5X18MM ONYX FRONTIER STENT WAS SUCCESSFULLY DEPLOYED ACROSS THE LESION. THE DELIVERY OF THE 2.5X26MM ONYX FRONTIER STENT WAS NOT PROVIDED ON THE IMAGES; HOWEVER, A DISLODGED STENT CAN CLEARLY BE SEEN IMMEDIATELY DISTAL TO THE TIP OF THE GUIDE CATHETER IN THE LM. THE DISLODGED STENT WAS SEQUENTIALLY DILATED WITH MULTIPLE BALLOONS RESULTING IN STENT DEPLOYMENT. A STENT WAS DEPLOYED DISTAL TO THE ORIGINALLY DEPLOYED STENT. THE OSTIUM OF THE LM TO LCX WAS BALLOONED WITH WHAT APPEARED TO BE THE 4.5 X 8MM NC EUPHORA BALLOON. THE PROCEDURAL IMAGES FAILED TO IDENTIFY THE CAUSE OF THE DISLODGEMENT AS IT WAS NOT CAPTURED ON THE IMAGES. A REVIEW OF THE IVUS IMAGES PROVIDED CONCLUDED THERE WAS MODERATE CALCIFICATION PRESENT. DOUBLE LAYERED STRUTS WERE SHOWN IN A FEW FRAMES. THERE WAS NO EVIDENCE OF SIGNIFICANT ANGIOGRAPHIC COMPLICATIONS. A 7F LAUNCHER GUIDE CATHETER WAS ENGAGED IN THE LEFT MAIN (LM). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20886272 | 2024 | PROMUS PREMIER | Unknown | ['Difficult to Remove', 'Reduced Blood Flow'] | E1 INITIAL REPORTER FACILITY NAME: (B)(6) HOSPITAL AFFILIATED TO (B)(6) UNIVERSITY. IT WAS REPORTED THAT THE PATIENT DIED. THE TARGET LESION WAS LOCATED IN THE LEFT ANTERIOR DESCENDING (LAD) ARTERY. AFTER LESION PREPARATION USING A BALLOON, A DISSECTION IN THE DIAGONAL BRANCH WAS NOTED. A PROMUS PREMIER DRUG-ELUTING STENT WAS ADVANCED TO TREAT THE DISSECTION. AFTER STENT DELIVERY TO THE DIAGONAL BRANCH, VARIOUS METHODS WERE PERFORMED IN ATTEMPTS TO WITHDRAWAL THE STENT BALLOON, UNSUCCESSFULLY. THE PATIENT EXPERIENCED SLOW BLOOD FLOW AND NO REFLOW. AFTER RESCUE, AN INTRA-AORTIC BALLOON PUMP (IABP) WAS APPLIED AND SURGICAL INTERVENTION WAS PERFORMED TO OPEN THE CHEST AND REMOVE THE DEVICE AND GUIDEWIRE. THE PATIENT WAS PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). TWELVE DAYS LATER, THE PATIENT DIED AFTER ECMO WITHDRAWAL. IT WAS REPORTED THAT THE PATIENT DIED. THE TARGET LESION WAS LOCATED IN THE LEFT ANTERIOR DESCENDING (LAD) ARTERY. AFTER LESION PREPARATION USING A BALLOON, A DISSECTION IN THE DIAGONAL BRANCH WAS NOTED. A PROMUS PREMIER DRUG-ELUTING STENT WAS ADVANCED TO TREAT THE DISSECTION. AFTER STENT DELIVERY TO THE DIAGONAL BRANCH, VARIOUS METHODS WERE PERFORMED IN ATTEMPTS TO WITHDRAWAL THE STENT BALLOON, UNSUCCESSFULLY. THE PATIENT EXPERIENCED SLOW BLOOD FLOW AND NO REFLOW. AFTER RESCUE, AN INTRA-AORTIC BALLOON PUMP (IABP) WAS APPLIED AND SURGICAL INTERVENTION WAS PERFORMED TO OPEN THE CHEST AND REMOVE THE DEVICE AND GUIDEWIRE. THE PATIENT WAS PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). TWELVE DAYS LATER, THE PATIENT DIED AFTER ECMO WITHDRAWAL. UPDATED AWARE DATE. E1 INITIAL REPORTER FACILITY NAME: (B)(6) HOSPITAL AFFILIATED TO (B)(6) MEDICAL UNIVERSITY.
20890785 | 2024 | RESOLUTE ONYX RX | Unknown | ['Inflation Problem', 'Activation, Positioning or Separation Problem', 'No Clinical Signs, Symptoms or Conditions'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING ONE RESOLUTE ONYX CORONARY DRUG ELUTING STENT, THE DISTAL PART OF THE BALLOON WOULD NOT INFLATE THE STENT BUT THE PROXIMAL PART DID INFLATE THE STENT. THE DEVICE MALFUNCTIONED INSIDE THE PATIENT AND EVASIVE ACTION HAD TO BE TAKEN TO SAVE THE CASE. THE STENT WOULD NOT FULLY DEPLOY DUE TO THE BALLOON. THE DEVICE WAS INSPECTED PRIOR TO USE WITH NO ISSUES. NO FURTHER PATIENT INJURY REPORTED. ADDITIONAL INFORMATION: ANNEX B AND D CODES. CORRECTION: ANNEX E <(>&<)> F CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20891671 | 2024 | ONYX FRONTIER | Male | ['Activation, Positioning or Separation Problem', 'Positioning Problem', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. THE PATIENT PRESENTED TO HOSPITAL WITH A NON-ST-ELEVATION MYOCARDIAL INFARCTION (NSTEMI). IT WAS REPORTED THAT ONE 3.0X8MM ONYX FRONTIER CORONARY DRUG ELUTING STENT DISLODGED IN THE VESSEL. THE DISLODGEMENT OCCURRED DURING DELIVERY TO/AT LESION. THE CASE WAS DIFFICULT AND COMPLEX. THE STENT WAS BEING USED TO TREAT A MODERATELY TORTUOUS, MODERATELY CALCIFIED LESION WITH 90% STENOSIS IN THE PROXIMAL DIAGONAL BRANCH (D1). THERE WAS TIMI 3 FLOW. THE DEVICE WAS INSPECTED PRIOR TO USE WITH NO ISSUES. NEGATIVE PREP WAS NOT PERFORMED. THE LESION WAS PRE-DILATED. RESISTANCE WAS ENCOUNTERED WHEN ADVANCING THE DEVICE. EXCESSIVE FORCE WAS NOT USED. A 6F GUIDE CATHETER CANNULATED THE LEFT MAIN (LM) AND TWO NON-MEDTRONIC (MDT) GUIDEWIRES CROSSED THE D1 LESION. A 2.5MM COMPLIANT BALLOON PRE-DILATED THE PROXIMAL D1 STENOSIS TO 8 ATM. INTRAVASCULAR ULTRASOUND (IVUS) WAS USED WHICH DEMONSTRATED DISEASE ALL THE WAY TO THE OSTIUM. A 2.0X26MM ONYX FRONTIER STENT WAS DEPLOYED FROM THE PROXIMAL TO MID D1 AT 12 ATM. FOLLOWING THIS AN ATTEMPT WAS MADE TO ADVANCE THE 3.0X8MM ONYX FRONTIER STENT TO THE OSTIUM OF THE D1, OVERLAPPING THE PROXIMAL EDGE OF THE INITIAL STENT. HOWEVER, THERE WAS SIGNIFICANT DIFFICULTY ADVANCING THE STENT BEYOND THE OSTIAL PORTION OF THE D1, AND DURING THIS PROCESS THE STENT STRIPPED OFF THE BALLOON WITHOUT DEPLOYMENT. AN ATTEMPT WAS MADE TO USE THE BALLOON TO RE-CAPTURE THE DISLODGED STENT INTO THE GUIDE, BUT THIS WAS NOT SUCCESSFUL. A LEFT ANTERIOR DESCENDING (LAD) ARTERY WIRE WAS ATTEMPTED TO WRAP AROUND THE STENT TO PULL IT OUT, BUT THIS WAS ALSO UNSUCCESSFUL. GUIDE AND WIRE ACCESS WERE THEN LOST, AND THE STENT WAS FREE FLOATING IN THE PROXIMAL LAD, LEADING INTO THE OSTIAL D1. A NEW NON-MDT GUIDEWIRE WAS USED AND A 4-8MM NON-MDT SNARE WAS USED TO SUCCESSFULLY SNARE THE DISLODGED STENT. THEN, ANOTHER 3.0X8MM ONYX FRONTIER STENT WAS ADVANCED TO THE D1, HOWEVER, THERE WAS DIFFICULTY CROSSING INTO THE PROXIMAL D1. THERE WAS CONCERN OF A POSSIBLE CORONARY DISSECTION AT THE OSTIUM OF THE D1 DUE TO THE PREVIOUSLY DEPLOYED STENT. EVENTUALLY, IT WAS POSSIBLE TO ADVANCE A 2.5MM NC BALLOON TO THE MID D1 FOR POST DILATION OF THE 2.0X26MM ONYX FRONTIER STENT, AS WELL AS PRE-DILATION OF THE OSTIUM D1 TO A MAXIMUM OF 18 ATM. THE SECOND 3.0X8MM ONYX FRONTIER STENT WAS THEN SUCCESSFULLY ADVANCED TO THE OSTIAL D1 OVERLAPPING THE PROXIMAL EDGE OF THE 2.0X26MM ONYX FRONTIER STENT AND WAS SUCCESSFULLY DEPLOYED TO 14 ATM. IVUS AFTER STENT DEPLOYMENT DEMONSTRATED A WELL-APPOSED STENT, HOWEVER, THERE WAS MILD UNDER-EXPANSION OF THE 3.0X8MM ONYX FRONTIER STENT. A 3.0MM NC BALLOON WAS THEN USED TO POST DILATE THE 3.0X8MM ONYX FRONTIER STENT TO 20 ATM. IVUS SHOWED GOOD STENT EXPANSION AND APPOSITION, WITH NO EVIDENCE OF STENT EDGE DISSECTION. POST INTERVENTION THERE WAS TIMI FLOW 3, AND 0% RESIDUAL STENOSIS. THE PATIENT IS ALIVE WITH NO FURTHER INJURY. ADDITIONAL INFORMATION: IT WAS NOT DIFFICULT TO REMOVE THE STYLET OR PROTECTIVE SHEATH BEFORE USE. THERE ARE NO COMPLAINTS AGAINST THE OTHER MEDTRONIC DEVICES USED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: ANNEX D CODE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20894761 | 2024 | SYNERGY XD | Male | ['Break', 'Difficult to Remove', 'Foreign Body In Patient'] | IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT PRESENTED FOR MULTIVESSEL PERCUTANEOUS CORONARY INTERVENTION WITH A NON-BOSTON SCIENTIFIC (BSC) HEART PUMP IN USE. THE 80 TO 90% STENOSED TARGET LESION WAS LOCATED IN THE SEVERELY CALCIFIED PROXIMAL CIRCUMFLEX ARTERY. A ROTAPRO AND A NON-BSC LITHOTRIPSY BALLOON WERE UTILIZED WITH NO PATIENT ISSUES. SEVEN STENTS WERE DEPLOYED WITH NO PATIENT COMPLICATIONS. A 3.50 X 24 MM SYNERGY XD DRUG-ELUTING STENT (DES) WAS DEPLOYED, AND AFTERWARD, THE PHYSICIAN INSTRUCTED THE TECHNICIAN TO INFLATE THE STENT DELIVERY SYSTEM BALLOON TO 30 ATM BEFORE REMOVING THE SYSTEM. THE BALLOON WAS INFLATED FOR 20 SECONDS AND THEN FULLY DEFLATED AFTER ABOUT 5 SECONDS. DURING REMOVAL, THE HYPOTUBE OR BALLOON GOT STUCK. RESISTANCE WAS ENCOUNTERED AND THE DEVICE WAS REMOVED WITH FORCE CAUSING THE HYPOTUBE TO SNAP INSIDE THE PATIENT. BALLOON MARKERBANDS SHOWED THAT A PORTION OF THE HYPOTUBE WAS SITTING BETWEEN THE LEFT MAIN AND AORTA. FOLLOWING CONSULTATION WITH A SURGEON, THE DECISION WAS MADE TO ATTEMPT TO RETRIEVE THE DEVICE FRAGMENT WITH A SNARE. ANESTHESIA WAS CALLED TO INTUBATE THE PATIENT. MULTIPLE UNSUCCESSFUL ATTEMPTS WERE MADE TO SNARE THE FRAGMENT BEFORE THE DECISION WAS MADE TO DEPLOY A STENT IN THE LEFT MAIN TO TACK THE FRAGMENT TO THE LEFT MAIN WALL AND COMPLETE THE PROCEDURE. THE PATIENT WAS ADMITTED TO THE INTENSIVE CARE UNIT FOR OBSERVATION. THE NEXT DAY THE PATIENT WAS BROUGHT BACK TO THE CATHETER LAB TO REAPPLY THE NON-BSC HEART PUMP, BUT THE PATIENT DIED.
20896444 | 2024 | XIENCE PROS EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Male | ['Break', 'Angina'] | IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT THE CIRCUMFLEX CORONARY ARTERY TO BYPASS GRAFT ANASTOMOSIS. THE 2.25X38 MM XIENCE PROS STENT WAS IMPLANTED ON 10/05/2023. THE PATIENT WAS HOSPITALIZED FOR CHEST PAIN ABOUT TEN MONTHS LATER, AND DURING A CONTROL ANGIOGRAPHY, THE STENT WAS FOUND TO BE BROKEN. NO TREATMENT WAS PROVIDED. NO ADDITIONAL INFORMATION WAS PROVIDED. MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B3: DATE OF EVENT 8/05/2024 IS ESTIMATED. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. A CINE WAS RECEIVED AND REVIEWED, THE REVIEWER CONCLUDED THAT THE CINE IMAGES PROVIDED WERE VERY GRAINY; HOWEVER, A POSSIBLE FRACTURE OF THE STENT COULD BE PARTIALLY VISUALIZED. THIS WAS AN INCIDENTAL FINDING AS THE CRITICAL LAD STENOSIS WAS DETERMINED TO BE THE CAUSE OF THE PATIENTÂ¿S CHEST PAIN, PER DR. TEYMEN. A PROBABLE CAUSE COULD NOT BE DETERMINED FROM THE IMAGES PROVIDED. THE REPORTED PATIENT EFFECT OF ANGINA IS LISTED IN THE XIENCE SIERRA EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. THE INVESTIGATION WAS UNABLE TO DETERMINE A CONCLUSIVE CAUSE FOR THE REPORTED BREAK OR PATIENT EFFECT. IT WAS REPORTED THAT THE STENT DELIVERY SYSTEM (SDS) WAS IMPLANTED HOWEVER THE PATIENT WAS HOSPITALIZED FOR CHEST PAIN ABOUT TEN MONTHS LATER, AND DURING A CONTROL ANGIOGRAPHY, THE STENT WAS FOUND TO BE BROKEN. FACTORS THAT MAY CONTRIBUTE TO STENT BREAK INCLUDE, BUT ARE NOT LIMITED TO, TENSILE STRENGTH, CORROSION, OVER INFLATION OF THE STENT, DEPLOYMENT TECHNIQUE, INTERACTION WITH ACCESSORY DEVICES, AND/OR INTERACTION WITH CHALLENGING ANATOMY. BASED ON THE INFORMATION PROVIDED AND BECAUSE THE DEVICE WAS NOT RETURNED FOR ANALYSIS, IT IS UNKNOWN WHAT MAY HAVE CAUSED THE REPORTED BREAK. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE.
20900021 | 2024 | ONYX FRONTIER | Male | ['Fracture', 'Activation, Positioning or Separation Problem', 'Material Deformation', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. IT WAS REPORTED THAT ONE ONYX FRONTIER CORONARY DRUG ELUTING STENT WAS IMPLANTED TO TREAT A NON-CALCIFIED, NON-TORTUOUS LESION WITH 80% STENOSIS IN THE OSTIUM-PROXIMAL CIRCUMFLEX (CX) ARTERY, WHEN STENT DEFORMATION OCCURRED IN VIVO POST-DEPLOYMENT. THE DEVICE WAS INSPECTED PRIOR TO USE WITH NO ISSUES. NEGATIVE PREP WAS PERFORMED WITH NO ISSUES. THE LESION WAS PRE-DILATED. THE DEVICE DID NOT PASS THROUGH A PREVIOUSLY DEPLOYED STENT. RESISTANCE WAS NOT ENCOUNTERED WHEN ADVANCING THE DEVICE. EXCESSIVE FORCE WAS NOT USED. IT WAS DETAILED THAT AFTER INFLATING THE STENT, THE STENT APPEARED FRACTURED FROM ALMOST THE MIDDLE, WHICH CAUSED THE STENT TO ELONGATE TOWARDS THE OSTIUM. IT WAS THEN DECIDED TO USE A STENT-IN-STENT APPROACH AND INFLATE THE PROXIMAL PART TO ENSURE THE OSTIUM WAS NOT COMPROMISED. THE PATIENT IS ALIVE WITH NO FURTHER INJURY. ADDITIONAL INFORMATION: THERE IS NO COMPLAINT AGAINST THE NC BALLOON USED, AND IT WORKED PERFECTLY. CORRECTION: ANNEX A CODE ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: THE LESION WAS MILD-MODERATELY CALCIFIED. THERE WERE NO PREVIOUS STENTS PRESENT AT THE SITE OF TREATMENT. IT WAS DETAILED THAT AFTER INFLATING THE STENT, THE STENT APPEARED FRACTURED FROM ALMOST THE MIDDLE, WHICH CAUSED THE STENT TO ELONGATE OUTSIDE THE OSTIUM, HANGING INSIDE THE AORTIC ROOT. THERE HAD BEEN NO DIFFICULTY DURING STENT DEPLOYMENT. THE DEVICE HAD BEEN INFLATED TO 14-16 ATMS. THE BALLOON OF THE ONYX FRONTIER DEVICE DID SUCCESSFULLY DEFLATE POST STENT DEPLOYMENT. THERE WAS NO DIFFICULTY NOTED DURING THE REMOVAL OF THE GUIDEWIRE. THE BALLOON OR DELIVERY SYSTEM DID NOT SNAG ON THE DEPLOYED STENT. IT WAS THEN DECIDED TO USE A STENT-IN-STENT APPROACH AND INFLATE THE PROXIMAL PART TO ENSURE THE OSTIUM WAS NOT COMPROMISED, AND TO TO PROTECT THE CRUSHED AREA OF THE STENT. A 5.0X12MM NC EUPHORA BALLOON WAS USED TO POST DILATE THE STENT TO 16-18 ATM, AS WELL AS A 4.0X12MM BALLOON TO 16-18 ATM. STENT DEFORMATION OCCURRED BEFORE USING THE NC BALLOONS. A4:PATIENT WEIGHT PROVIDED B7:RELEVANT HISTORY MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. PROCEDURAL IMAGE ANALYSIS: FLUOROSCOPY IMAGES WERE PROVIDED FOR REVIEW. THE PATIENT HAD PREVIOUSLY UNDERGONE A CABG PROCEDURE, AND IT APPEARS THAT THE LCX WAS BY PASSED DURING THAT PROCEDURE. THE LCX WAS OCCLUDED, AND A PROCEDURAL WIRE WAS SUCCESSFULLY DELIVERED INTO THE VESSEL. THE TARGET LESION WAS IMMEDIATELY DISTAL TO THE SITE OF THE OCCLUSION. THE VESSEL WAS NOT TORTUOUS, BUT THE LESION WAS ON A BEND IN THE VESSEL. A LONG STENT, MOST LIKELY THE ONYX FRONTIER STENT WAS POSITIONED OVERLAPPING WITH A PREVIOUSLY DEPLOYED STENT AND WAS SUCCESSFULLY DEPLOYED. THE AREA OF THE BEND IN THE VESSEL WAS AT THE MIDDLE OF THE STENT AND THE STENT APPEARED TO CONFORM TO THE VESSEL AT THIS POINT RESULTING IN THE PREVIOUSLY ADJACENT STENT NON-WELDED WIRE SEGMENT TO SEPARATE. THIS GAVE THE IMPRESSION OF A FRACTURE, BUT SINCE NOT ALL POINTS ON THE ONYX FRONTIER STENT ARE WELDED THERE MOST LIKELY NO FRACTURE BUT DEFORMATION OF THE STENT STRUCTURE DUE TO THE VESSEL PROFILE. AS STATED IN THE INFORMATION PROVIDED FROM THE ACCOUNT THE STENT AND THE SUBSEQUENT STENTS THAT WERE DEPLOYED WERE POST DILATED IN ORDER TO OBTAIN AN ACCEPTABLE RESULT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20900516 | 2024 | ONYX FRONTIER | Male | ['Difficult to Remove', 'Device Dislodged or Dislocated', 'Device Embedded In Tissue or Plaque', 'Insufficient Information'] | IT WAS REPORTED THAT ONE ONYX FRONTIER CORONARY DRUG ELUTING STENT DISLODGED DURING DELIVERY TO/AT THE LESION. THE STENT WAS UNABLE TO BE REMOVED FROM THE PATIENT AND ANOTHER STENT WAS USED TO PUSH THE DISLODGED STENT AGAINST THE VESSEL WALL. THE LESION BEING TREATED WAS SEVERELY CALCIFIED AND LOCATED IN THE RIGHT CORONARY ARTERY (RCA). THE DEVICE WAS INSPECTED PRIOR TO USE WITH NO ISSUES. THE DEVICE DID NOT PASS THROUGH A PREVIOUSLY DEPLOYED STENT. RESISTANCE WAS ENCOUNTERED WHILE ADVANCING THE DEVICE. THE PATIENT IS ALIVE WITH NO FURTHER INJURY. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: THE DEVICE WAS PREPPED PER IFU WITH NO ISSUES NOTED. THE LESION WAS PRE-DILATED. EXCESSIVE FORCE WAS USED DURING DELIVERY. PATIENT WEIGHT PROVIDED. EVENT DATE PROVIDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: ANNEX D CODE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20901311 | 2024 | ONYX FRONTIER | Male | ['Deformation Due to Compressive Stress', 'Device Damaged by Another Device', 'Device Dislodged or Dislocated', 'Device Embedded In Tissue or Plaque', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. IT WAS REPORTED THAT DURING THE ONE ONYX FRONTIER CORONARY DRUG ELUTING STENT DISLODGED DURING DELIVERY TO/AT THE LESION. THE STENT DISLODGED WHEN USED WITH A NON-MEDTRONIC GUIDE EXTENSION CATHETER. THE DISLODGED STENT WAS NOT REMOVED. A SMALL BALLOON WAS INSERTED THROUGH THE DISLODGED STENT TO DEPLOY THE STENT. THE DEVICE WAS INSPECTED BEFORE USE WITH NO ISSUES. NEGATIVE PREP WAS PERFORMED WITH NO ISSUES. EXCESSIVE FORCE WAS NOT USED DURING DELIVERY. THE PATIENT IS ALIVE WITH NO FURTHER INJURY. ADDITIONAL INFORMATION: LOT NUMBER UPDATED. PRODUCT ANALYSIS: THE STENT WAS NOT PRESENT ON THE BALLOON AND DID NOT RETURN FOR ANALYSIS. A KINK WAS EVIDENT TO THE HYPOTUBE. THE BALLOON FOLDS WERE PARTIALLY EXPANDED. CRIMP IMPRESSIONS WERE VISIBLE ON THE EXPOSED BALLOON SURFACE. NO OTHER DAMAGE EVIDENT TO THE REMAINDER OF THE DEVICE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20915123 | 2024 | PROMUS ELITE | Female | ['Difficult to Remove', 'Failure to Advance', 'Device Dislodged or Dislocated', 'Material Deformation', 'Foreign Body In Patient'] | IT WAS REPORTED THAT REMOVAL DIFFICULTY OCCURRED. THE PATIENT PRESENTED WITH MYOCARDIAL INFARCTION AND UNDERWENT PERCUTANEOUS TRANSLUMINAL CORONARY ANGIOPLASTY. VASCULAR ACCESS WAS OBTAINED VIA THE RADIAL ARTERY. THE 80% STENOSED TARGET LESION WITH A BEND OF ALMOST 90 DEGREES WAS LOCATED IN THE SEVERELY CALCIFIED PROXIMAL LEFT CIRCUMFLEX ARTERY. FOLLOWING PRE-DILATATION, A 28 X 3.00 PROMUS ELITE DRUG-ELUTING STENT WAS ADVANCED FOR TREATMENT; HOWEVER, THE STENT FAILED TO CROSS THE LESION DUE TO THE SEVERE CALCIFICATION AND WHEN FORCED WAS APPLIED, THE STENT STRUTS GOT DEFORMED. THE DECISION WAS MADE TO CANCEL THE PROCEDURE AND WHEN THE DEVICE WAS BEING SNARED TO REMOVE, THE STENT GOT STUCK AND REMAINED IN THE RADIAL ARTERY. THERE WERE NO PATIENT COMPLICATIONS.
20921521 | 2024 | XIENCE XPEDITION 48 EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Unknown | ['Patient-Device Incompatibility', 'Vascular Dissection'] | IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT THE PROXIMAL LEFT ANTERIOR DESCENDING (LAD) CORONARY ARTERY WITH HEAVY TORTUOSITY. THE 3.0X48 MM XIENCE XPEDITION STENT WAS BEING IMPLANTED WHEN A DISSECTION WAS NOTED AND THE STENT WAS HANGING IN THE LUMEN DUE TO POOR WALL APPOSITION. ANOTHER 3.0X48 MM XIENCE XPEDITION STENT WAS USED TO TREAT THE DISSECTION AND TREAT THE HANGING STENT. THERE WERE NO ADVERSE PATIENT SEQUELA. NO ADDITIONAL INFORMATION WAS PROVIDED. MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. A CONCLUSIVE CAUSE FOR THE REPORTED DIFFICULTY AND PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. IT WAS REPORTED THAT THE STENT DELIVERY SYSTEM (SDS) WAS IMPLANTED AND WAS NOTED THAT THE STENT WAS NOT APPOSE TO THE VESSEL WALL. FACTORS THAT MAY CONTRIBUTE TO PATIENT-DEVICE INCOMPATIBILITY (STENT NOT FULLY APPOSED TO VESSEL WALL) INCLUDE, BUT ARE NOT LIMITED TO, STENT PLACEMENT, PATIENT ANATOMICAL MORPHOLOGY (STENOSED), INFLATION PROBLEM, OR UNDER SIZING OF THE VESSEL. BASED ON THE INFORMATION PROVIDED AND BECAUSE THE DEVICE WAS NOT RETURNED FOR ANALYSIS, IT IS UNKNOWN WHAT MAY HAVE CAUSED THE REPORTED PATIENT-DEVICE INCOMPATIBILITY. THE REPORTED PATIENT EFFECT OF DISSECTION IS LISTED IN THE XIENCE XPEDITION EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE.
20947435 | 2024 | SYNERGY | Male | ['Material Deformation', 'Vascular Dissection'] | IT WAS REPORTED THAT A DISSECTION OCCURRED. THE 80% STENOSED TARGET LESION WAS LOCATED IN THE MODERATELY TORTUOUS AND SEVERELY CALCIFIED LEFT ANTERIOR DESCENDING ARTERY (LAD). A 3.50 X 20 MM SYNERGY DRUG-ELUTING STENT WAS DEPLOYED. AFTER POST-DILATION, A SMALL PROXIMAL STENT EDGE DISSECTION WAS OBSERVED. STENT DAMAGE WAS ALSO NOTED. THE DISSECTION WAS COVERED WITH A 3.50 X 12 MM SYNERGY DRUG-ELUTING STENT AND THE PROCEDURE WAS COMPLETED. THE PATIENT WAS STABLE POST-PROCEDURE.
20951025 | 2024 | XIENCE SIERRAÃ‚Â¿ | Unknown | ['Difficult to Advance', 'Vascular Dissection'] | PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE REPORTED INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED DIFFICULT TO ADVANCE AND SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO CIRCUMSTANCES OF THE PROCEDURE; HOWEVER, A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. IT WAS REPORTED THAT THE STENT DELIVERY SYSTEM (SDS) MET RESISTANCE WITH HEAVY CALCIFICATION AND TORTUOSITY IN THE LESION, RESULTING IN DIFFICULT TO ADVANCE. THE REPORTED PATIENT EFFECT OF DISSECTION IS LISTED IN THE XIENCE SIERRA EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT THE LEFT CIRCUMFLEX CORONARY ARTERY WITH HEAVY CALCIFICATION AND HEAVY TORTUOSITY. THE 2.75X28 MM XIENCE SIERRA STENT DELIVERY SYSTEM (SDS) WAS ADVANCED TO THE LESION WITH DIFFICULTY NOTED WITH THE ANATOMY. THE STENT WAS IMPLANTED AND A DISSECTION OCCURRED. ANOTHER STENT WAS USED TO TREAT THE DISSECTION. THERE WERE NO ADVERSE PATIENT SEQUELA. NO ADDITIONAL INFORMATION WAS PROVIDED.
20951121 | 2024 | RESOLUTE INTEGRITY RX | Unknown | ['Material Deformation', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING ONE RESOLUTE INTEGRITY CORONARY DRUG ELUTING STENT, STENT DEFORMATION OCCURRED IN VIVO DURING POSITIONING/ADVANCEMENT. DESPITE MULTIPLE POST DILATIONS THE STENT RE COLLAPSED AND ANOTHER STENT HAD TO BE PLACED TO CRUSH THE COLLAPSED STENT AGAINST THE WALL. THE PROCEDURE WAS COMPLETED. THE DEVICE WAS INSPECTED WITH NO ISSUES NOTED. THE DEVICE DID NOT PASS THROUGH A PREVIOUSLY DEPLOYED STENT. RESISTANCE WAS NOT ENCOUNTERED WHEN ADVANCING THE DEVICE. EXCESSIVE FORCE WAS NOT USED DURING DELIVERY. THE PATIENT IS ALIVE WITH NO FURTHER INJURY. ADDITIONAL INFORMATION: THE STENT DEFORMATION/STENT COLLAPSE OCCURRED POST DEPLOYMENT OF THE STENT. NO INFLATION DIFFICULTIES WERE ENCOUNTERED DURING THE STENT DEPLOYMENT. THE STENT WAS NOT PARTIALLY DEPLOYED DURING THE INITIAL INFLATION. THERE WERE NO ISSUES NOTED WITH THE REPLACEMENT DEVICE. ANNEX D CODE ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20951744 | 2024 | ONYX FRONTIER | Male | ['Deflation Problem', 'Material Deformation', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. IT WAS REPORTED THAT THERE WERE BALLOON DEFLATION DIFFICULTIES WHEN USING ONE ONYX FRONTIER CORONARY DRUG ELUTING STENT. THE DEVICE REQUIRED ABOUT 180 SECONDS TO CONTRACT AT THE LESION SITE. THE DEVICE WAS BEING USED TO TREAT A LESION IN THE LEFT ANTERIOR DESCENDING (LAD) BRANCH (#6-#7). THE DEVICE WAS INSPECTED PRIOR TO USE WITH NO PROBLEMS OBSERVED. NEGATIVE PREP WAS PERFORMED WITH NO PROBLEMS OBSERVED. THE LESION WAS PRE-DILATED. THERE WAS NO RESISTANCE/DISCOMFORT WHEN DELIVERING THE DEVICE. EXCESSIVE FORCE WAS NOT APPLIED. AFTER STENT PLACEMENT BALLOON DEFLATION WAS DIFFICULT, TAKING ABOUT 3 MINUTES FOR AIR REMOVAL. IT WAS SUSPECTING THERE WAS A MALFUNCTION OF THE INDEFLATOR, THEREFORE, AIR REMOVAL WAS ATTEMPTED USING A SYRINGE. THE STENT WAS IMPLANTED. POST-PROCEDURE EXAMINATION OUTSIDE THE BODY SHOWED THAT EXPANSION WAS POSSIBLE, BUT DEFLATION REMAINED DIFFICULT. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. PRODUCT ANALYSIS: THE DEVICE WAS RETURNED FOR ANALYSIS. THE RETURNED DEVICE SHOWED EVIDENCE OF BALLOON INFLATION AND CONTRAST REMAINED IN THE BALLOON. THE INFLATION LUMEN/PROXIMAL BOND WAS NECKED. DURING INFLATION TESTING IT WAS POSSIBLE TO INFLATE THE BALLOON TO A NOMINAL OF 12ATM; LIQUID DID PASS THROUGH THE NECKED SECTION OF THE INFLATION LUMEN/PROXIMAL BOND. DURING DEFLATION TESTING LIQUID COULD NOT PASS BACK THROUGH THE NECKED SECTION AND THE BALLOON FAILED TO DEFLATE. ADDITIONAL INFORMATION: IT WAS NOT DIFFICULT TO REMOVE THE PROTECTIVE SHEATH OR THE PACKAGING STYLETTE. THE DEVICE DID NOT PASS THROUGH A PREVIOUSLY PLACED STENT. THE BALLOON WAS INFLATED TO 12 ATM FOR 5 SECONDS. THERE WERE NO DIFFICULTIES NOTED DURING INFLATION OF THE DEVICE. DEFLATION DIFFICULTIES WERE NOTED AFTER THE FIRST INFLATION. THE DEVICE WAS NOT MOVED OR REPOSITIONED WHILE INFLATED. A CONCENTRATION OF 2:1 CONTRAST / SALINE WAS USED. A BALLOON WAS USED FOR STENT POST-DILATION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20953521 | 2024 | XIENCE PROA EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Unknown | ['Difficult to Remove', 'Difficult to Advance', 'Device Dislodged or Dislocated', 'No Clinical Signs, Symptoms or Conditions'] | MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. IT WAS REPORTED THAT 2.25X18MM XIENCE PROA STENT DELIVERY SYSTEM WAS ATTEMPTED BE ADVANCED THROUGH THE CATHETER; HOWEVER, THE STENT BECAME LOOSE [DISLODGED] AND REMAINED STUCK WITH THE CATHETER. THERE WAS NO ADVERSE PATIENT EFFECT AND NO CLINICALLY SIGNIFICANT DELAY IN THE PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED. SUBSEQUENTLY, AFTER THE INITIAL REPORT WAS FILED IT WAS NOTED THAT 2.25X18MM XIENCE PROA STENT DELIVERY SYSTEM WAS ATTEMPTED TO BE ADVANCED HOWEVER, THE STENT BECAME STUCK WITH A PREVIOUSLY IMPLANTED STENT AND DISLODGED. THE DISLODGED STENT WAS RETRIEVED VIA SNARE. ANOTHER DEVICE WAS USED TO SUCCESSFULLY COMPLETE THE PROCEDURE. THERE WAS NO ADVERSE PATIENT SEQUELA AND NO CLINICALLY SIGNIFICANT DELAY IN THE PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED. MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. H1 CORRECTION: UPDATED FROM MALFUNCTION TO SERIOUS INJURY H6 CORRECTION: CODE 2199 REMOVED. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. THE INVESTIGATION WAS UNABLE TO DETERMINE A CONCLUSIVE CAUSE FOR THE REPORTED DIFFICULTIES; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. IN THIS CASE, IT IS LIKELY THAT THE MANIPULATION OF THE STENT DELIVERY SYSTEM (SDS) DURING INTERACTION WITH PREVIOUS IMPLANTED STENT CONTRIBUTED TO THE REPORTED STENT DISLODGEMENT OF THE STENT IN THE LESION. ADDITIONALLY, IT WAS REPORTED THAT THE DISLODGED STENT WAS RETRIEVED VIA SNARE FROM PATIENTÂ¿S LESION. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING.
20955682 | 2024 | XIENCE XPEDITION 48 EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Unknown | ['Break', 'No Clinical Signs, Symptoms or Conditions'] | MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT A RIGHT CORONARY ARTERY. THE 3.0X48MM XIENCE XPEDITION WAS IMPLANTED AT 11 TO 12 ATMOSPHERES; HOWEVER, THE STENT STRUT WAS NOTED TO BE FRACTURED. THEREFORE, A 3.0X23MM XIENCE ALPINE STENT WAS IMPLANTED TO COVERED THE FRACTURED STENT. THERE WAS NO ADVERSE PATIENT SEQUELA AND NO CLINICALLY SIGNIFICANT DELAY IN THE PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. A CONCLUSIVE CAUSE FOR THE REPORTED BREAK CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. FACTORS THAT MAY CONTRIBUTE TO STENT BREAK INCLUDE, BUT ARE NOT LIMITED TO, INFLATION ABOVE RATED BURST PRESSURE, INTERACTION WITH CHALLENGING ANATOMY, INTERACTION WITH ACCESSORY DEVICES OR MANIPULATION AGAINST RESISTANCE. BASED ON THE INFORMATION PROVIDED AND BECAUSE THE DEVICE WAS NOT RETURNED FOR ANALYSIS, IT IS UNKNOWN WHAT MAY HAVE CAUSED THE REPORTED BREAK. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING.
20961443 | 2024 | ONYX FRONTIER | Female | ['Deflation Problem', 'Entrapment of Device', 'Difficult to Remove', 'Deformation Due to Compressive Stress', 'Detachment of Device or Device Component', 'Material Deformation', 'Positioning Problem', 'Device Embedded In Tissue or Plaque', 'Insufficient Information'] | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. IT WAS REPORTED THAT ONE ONYX FRONTIER CORONARY DRUG ELUTING STENT ENCOUNTERED BALLOONÂ DEFLATION DIFFICULTIES AND THAT THE DEVICE WOULD NOT DEFLATE AT THE LESION SITE. THE BALLOON WAS DIFFICULT TO REMOVE FOLLOWING STENT DEPLOYMENT, AND FORCEFUL REMOVAL OF THE DELIVERY SYSTEM CAUSED THE BALLOON TO SEPARATE AND DETACH FROM THE DELIVERY SYSTEM AND REMAIN IN PLACE IN THE PATIENT. THE DELIVERY OF THE 3.5X12MM ONYX FRONTIER STENTÂ WAS DIFFICULT ANDÂ THERE WERE MULTIPLE KINKS ALONG THE DELIVERY SYSTEM. THE STENT WAS SUCCESSFULLY DEPLOYED, HOWEVER THE BALLOON STRIPPED OFF THE DELIVERY SYSTEM UPON REMOVAL. THE PATIENT WAS SENT TO SURGERYÂ FOR REMOVAL OF THE BALLOON ANDÂ IS NOW RECOVERING MAKING SLOW PROGRESS. THE DEVICE WAS INSPECTED WITH NO ISSUES. NEGATIVE PREP WAS PERFORMED WITH NO ISSUES. THE LESION WAS PRE-DILATED WITH A NON-MEDTRONIC DEVICE. THE DEVICE DID PASS THROUGH A PREVIOUSLY DEPLOYED STENT. RESISTANCE WAS ENCOUNTERED WHEN ADVANCING THE DEVICE. EXCESSIVE FORCE WAS USED DURING DELIVERY. THE LESION WAS IN THE MIDÂ RIGHT CORONARY ARTERY (RCA) AND HAD MILD TORTUOSITY, SEVERE CALCIFICATION WITH 80% STENOSIS. ONEÂ ONYX FRONTIER CORONARY DRUG ELUTING STENTÂ (3.5X8MM) WAS DEPLOYED SUCCESSFULLY PRIOR TO THE ISSUES ENCOUNTERED. THE PATIENT IS ALIVE WITH INJURY. ADDITIONAL INFORMATION: THE BALLOON FAILED TO DEFLATE. IT WAS NOT DIFFICULT TO REMOVE THE PROTECTIVE SHEATH OR THE PACKAGING STYLETTE. THE LESION WAS PRE-DILATED WITH A NON-MEDTRONIC CUTTING BALLOON, AND SHOCKWAVE WAS PERFORMED. DIFFICULTIES WERE NOT NOTED DURING INFLATION OF THE DEVICE. DEFLATION DIFFICULTIES WERE NOTED AFTER THE FIRST INFLATION. A 50/50 CONCENTRATION OF CONTRAST / SALINE WAS USED. THE DEVICE WAS NOT KINKED AND RE-STRAIGHTENED DURING USE. THE STENT WAS SUFFICIENTLY EXPANDED PRIOR TO ATTEMPTING REMOVAL OF THE BALLOON, AND SUFFICIENT TIME WAS GIVEN TO ALLOW THE BALLOON TO FULLY DEFLATE PRIOR TO ATTEMPTING REMOVAL. THERE WAS NO DAMAGE NOTED TO THE GUIDEWIRE ON REMOVAL FROM THE PATIENT. THE PATIENT WAS SENT TO SURGERY FOR REMOVAL OF THE BALLOON VIA OPEN HEART SURGERY. THE PATIENT IS HEMODYNAMICALLY STABLE WITH NO CHEST PAIN AND NO ECG CHANGES. PATIENT WEIGHT. CORRECTION: THE 3.5X12MM ONYX FRONTIER DES WAS DIFFICULT TO DELIVER AND A DECISION WAS MADE TO REMOVE AND USE A SHORTER STENT FIRST. ONE 3.5X8MM ONYX FRONTIER DES WAS DEPLOYED SUCCESSFULLY PRIOR TO THE ISSUES ENCOUNTERED. THE 3.5X12MM ONYX FRONTIER DES WAS RE-INSERTED. THE STENT WAS SUCCESSFULLY DEPLOYED; HOWEVER, IT WAS NOT POSSIBLE TO DEFLATE THE STENT BALLOON, AND IT WAS AGGRESSIVELY TRIED TO REMOVE DELIVERY SYSTEM. EVENTUALLY THIS WAS SUCCESSFUL, BUT THE BALLOON STRIPPED OFF THE DELIVERY SYSTEM UPON REMOVAL AND REMAINED IN THE PROXIMAL RIGHT CORONARY ARTERY (RCA). THE DEVICE WAS NOT MOVED OR REPOSITIONED WHILE INFLATED. THE SURGEON NOTED THAT TO REMOVE THE BALLOON THEY HAD TO PUSH FORWARD THEN PULL BACK. ANNEX E CODE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. ADDITIONAL INFORMATION: A PRESSURE OF 12 ATM WAS APPLIED DURING BALLOON INFLATION. IT IS BELIEVED THAT THAT THE STENT WAS DEFORMED DURING BALLOON REMOVAL AND/OR DURING PREVIOUS ATTEMPTS TO DELIVER. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
20964330 | 2024 | XIENCE SKYPOINTÃ‚Â¿ | Female | ['Device Dislodged or Dislocated', 'Material Deformation', 'Foreign Body In Patient'] | MEDWATCH MW5162648 RECEIVED WHICH STATES: THE STENT WAS PREPARED TO LOAD ONTO THE WIRE AND VISUALIZED. WHEN THE STENT GOT THROUGH THE CATHETER TO THE CORONARY THE STENT WAS VISIBLY DAMAGED. THE PHYSICIAN PULLED THE STENT DELIVERY SYSTEM OUT THE OF THE SYSTEM AND NOTED THE STENT WAS NOT ON THE SYSTEM. STENT WAS NOT IDENTIFIED. PHYSICIAN REMOVED PRODUCT, EXCEPT THE SHEATH. HE INSERTED A DIAGNOSTIC CATHETER TO TAKE FLURO GUIDED IMAGES OF THE UPPER BODY (EXCLUDING HEAD AND NECK) AND LOWER EXTREMITIES. NO STENT WAS VISUALIZED. PRODUCT RETAINED AND KEPT FOR REVIEW. PHYSICIAN ABORTED INTERVENING ON THAT VESSEL. CASE WAS FINISHED AND PATIENT WAS TAKEN TO POST-OP AREA. IMAGING ORDERS PLACED TO EVALUATE FOR RETAINED OBJECT. MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. A CONCLUSIVE CAUSE FOR THE REPORTED DIFFICULTIES AND PATIENT EFFECTS, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. IT WAS REPORTED THAT THE STENT DELIVERY SYSTEM'S (SDS) STENT WAS OBSERVED TO BE DEFORMED AFTER ADVANCEMENT. THE PHYSICIAN REMOVED THE SDS AND NOTED THE STENT WAS NOT ON THE SYSTEM. FACTORS THAT MAY CONTRIBUTE TO MATERIAL DEFORMATION OR STENT DISLODGEMENT INCLUDE, BUT ARE NOT LIMITED TO, INTERACTION WITH ACCESSORY DEVICES, PRODUCT DAMAGE DURING HANDLING/ PREPARATION OF DEVICE, OR INTERACTION WITH A CHALLENGING LESION. BASED ON THE INFORMATION PROVIDED AND BECAUSE THE DEVICE WAS NOT RETURNED FOR ANALYSIS, IT IS UNKNOWN WHAT MAY HAVE CAUSED THE REPORTED MATERIAL DEFORMATION OR STENT DISLODGEMENT. THE DISLODGED STENT REMAINS WITHIN THE ANATOMY. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE.
20964598 | 2024 | XIENCE SKYPOINTÃ‚Â¿ | Unknown | ['Failure to Advance', 'Difficult to Advance', 'Device Dislodged or Dislocated', 'Material Deformation', 'No Clinical Signs, Symptoms or Conditions'] | IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT THE RIGHT CORONARY ARTERY WITH HEAVY CALCIFICATION, MODERATE TORTUOSITY AND 75% STENOSIS. THE 3.5X48 MM XIENCE SKYPOINT STENT DELIVERY SYSTEM (SDS) FAILED TO CROSS THE LESION DUE TO THE ANATOMY AND THE STENT DISLODGED FROM THE DELIVERY SYSTEM. THE STENT STRUTS BECAME DAMAGED AND STUCK IN THE INTRODUCER SHEATH DURING ADVANCEMENT. THE STENT WAS STUCK IN THE FEMORAL ARTERY AND WAS RETRIEVED BY THE VASCULAR SURGEON VIA VASCULAR CUT DOWN. ANOTHER XIENCE STENT WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT SEQUELA. NO ADDITIONAL INFORMATION WAS PROVIDED. MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED ISSUE AND TREATMENT APPEARS TO BE RELATED TO OPERATIONAL CONTEXT. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. H6: MEDICAL DEVICE PROBLEM CODE 2524 WAS REMOVED AND 2920 ADDED.
20964661 | 2024 | XIENCE SKYPOINTÃ‚Â¿ | Male | ['Leak/Splash', 'Failure to Deflate', 'Non specific EKG/ECG Changes'] | MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT THE LEFT ANTERIOR DESCENDING (LAD) CORONARY ARTERY. OPTICAL COHERENCE TOMOGRAPHY (OCT) WAS PERFORMED AND THEN ATHERECTOMY WAS PERFORMED WITH A DEVICE FROM ANOTHER MANUFACTURER. AFTER ATHERECTOMY WAS COMPLETED, ANOTHER OCT RUN WAS PERFORMED. A 2.75X15 MM NC TREK NEO BALLOON DILATATION CATHETER (BDC) WAS ADVANCED TO THE LESION AND INFLATED FOR PRE-DILATATION AND THEN THE 3.0X38 MM XIENCE SKYPOINT STENT DELIVERY SYSTEM (SDS) WAS ADVANCED TO THE LESION AND THE STENT WAS DEPLOYED SUCCESSFULLY. AFTER DEPLOYMENT, THE SDS BALLOON WAS PULLED BACK TO THE PROXIMAL LAD TO USE THE MARKERS TO GAUGE THE LENGTH FOR A SECOND STENT. THE BALLOON WAS THEN ADVANCED BACK INTO THE STENT AND POST DILATATION WAS PERFORMED. AFTER INFLATION, THE BALLOON WOULD NOT DEFLATE. A SECOND INDEFLATOR WAS USED BUT THIS ALSO DID NOT DEFLATE THE BALLOON. THE PATIENT BEGAN TO HAVE ST RHYTHM CHANGES. A BALLOON PUMP WAS INSERTED AND ATTEMPTS WERE MADE TO DEFLATE THE BALLOON. THE BALLOON WAS PARTIALLY INFLATED IN THE LAD FOR ABOUT 17 MINUTES. EVENTUALLY THE BALLOON CAME DOWN AND THE BALLOON WAS TESTED OUTSIDE THE ANATOMY. IT APPEARED THERE WAS A LEAK IN THE SHAFT OF THE DELIVERY SYSTEM. THERE WERE NO REPORTED ADVERSE PATIENT SEQUELA. NO ADDITIONAL INFORMATION WAS PROVIDED. VISUAL INSPECTION AND FUNCTIONAL TESTING WERE PERFORMED ON THE RETURNED DEVICE. THE REPORTED LEAK WAS CONFIRMED. THE REPORTED FAILURE TO DEFLATE COULD NOT BE CONFIRMED DUE TO THE CONDITION THE DEVICE WAS RECEIVED FROM THE ACCOUNT. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE REPORTED INFORMATION AND THE OBSERVATIONS FROM THE RETURNED ANALYSIS, THE INVESTIGATION DETERMINED THAT THE REPORTED ISSUE AND TREATMENT APPEARS TO BE RELATED TO CIRCUMSTANCES OF THE PROCEDURE. BOTH THE OBSERVED INNER/OUTER MEMBER DAMAGE (PINCHED APPEARANCE) AND TORN GUIDE WIRE EXIT NOTCH OCCURRED POST DEPLOYMENT OF THE STENT. THE STENT WAS REPORTEDLY, DEPLOYED SUCCESSFULLY AND THE DELIVERY SYSTEM HAD NO ISSUES WITH INFLATION AND/OR DEFLATION OF THE DEVICE DURING STENT DEPLOYMENT. IT IS LIKELY THAT THE OBSERVED INDENT (PINCHED APPEARANCE) OF THE INNER AND OUTER MEMBER IDENTIFIED DURING RETURN ANALYSIS BLOCKED THE INFLATION LUMEN CAUSING THE REPORTED FAILURE TO DEFLATE AND PATIENT EFFECT OF EKG/ECG CHANGES. REPORTEDLY, AFTER DEPLOYMENT, THE STENT DELIVERY SYSTEM (SDS) BALLOON WAS PULLED BACK TO THE PROXIMAL LAD AND THEN ADVANCED BACK INTO THE DEPLOYED STENT. DURING THIS MOVEMENT OF THE DELIVERY SYSTEM, IT IS LIKELY THE OBSERVED INDENT (PINCHED APPEARANCE) TOOK PLACE AS THE DELIVERY SYSTEM WAS INFLATED AND DEFLATED WITHOUT ISSUE PRIOR. THE BALLOON WAS THEN ADVANCED BACK INTO THE STENT AND POST DILATATION WAS PERFORMED. AFTER INFLATION, THE BALLOON WOULD NOT DEFLATE AS THE INNER MEMBER WAS PINCHED. HANDLING AND/OR MANIPULATION OF THE DEVICE DURING THE DELAYED REMOVAL IN COMBINATION WITH INTERACTION WITH THE PROCEDURAL GUIDE WIRE LIKELY CAUSED THE OBSERVED TORN NOTCH AND REPORTED LEAK. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE.
20965604 | 2024 | XIENCE SKYPOINTÃ‚Â¿ | Female | ['Difficult to Advance', 'Hypoxia', 'Obstruction/Occlusion', 'Unspecified Tissue Injury'] | IT WAS REPORTED IN A RESEARCH ARTICLE CASE SUMMARY THAT THE PROCEDURE WAS TO TREAT A LONG TORTUOUS PATENT DUCTUS ARTERIOSIS (PDA) FROM THE UNDERSURFACE OF THE AORTA SUPPLYING CONFLUENT BUT MILDLY HYPOPLASTIC PULMONARY ARTERIES. A NON-ABBOTT .014 GUIDE WIRE WAS ADVANCED AND A 3.5X33 MM XIENCE STENT DELIVERY SYSTEM (SDS) WAS ATTEMPTED; HOWEVER THE DEVICE DID NOT CROSS DUE TO THE ANATOMY. AT THIS TIME THE PATIENT'S OXYGEN SATURATION DRIFTED TO 60%, SO A 3.5X23 MM XIENCE SKYPOINT SDS WAS QUICKLY ADVANCED AND DEPLOYED IN THE DISTAL PDA AND SATURATIONS IMPROVED FROM 60 TO 80%. IT WAS NOTED THAT THERE WAS STILL UNCOVERED LESION TOWARDS THE AORTA. THE 3.5X12 MM SDS WAS ATTEMPTED TO BE ADVANCED BUT HAD DIFFICULTY WITH CROSSING THE 3.5X23 MM XIENCE SKYPOINT STENT. THE SDS WAS NEGOTIATED INTO PLACE TO EXTEND THE PRIOR STENT PROXIMALLY AND WAS DEPLOYED. SATURATIONS IMMEDIATELY WORSENED AND AN OBSTRUCTION WAS NOTED WHICH WAS STATED TO BE AN INVAGINATION OF THE SOFT TISSUE OF THE PDA AND NOT THROMBUS. HEPARIN BOLUS WAS PROVIDED AND THE PATIENT WAS EMERGENTLY CANNULATED ONTO VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA-ECMO) AND BROUGHT TO THE OPERATING ROOM FOR EMERGENT SHUNT PLACEMENT. THE STENT WAS EXPLANTED AND THE PATIENT WAS DISCHARGED HOME 10 DAYS POST SURGERY. ADDITIONAL INFORMATION CAN BE FOUND IN THE ARTICLE "ACUTE PATENT DUCTUS ARTERIOSUS STENT OCCLUSION: A UNIQUE PERSPECTIVE". THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE PRODUCTION RECORDS AND THE CORRECTIVE AND PREVENTIVE ACTIONS (CAPA) DATABASE WAS PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED NO INDICATION OF A LOT SPECIFIC ISSUE. A CONCLUSIVE CAUSE FOR THE REPORTED DIFFICULT TO ADVANCE, UNEXPECTED MEDICAL INTERVENTION, SURGICAL INTERVENTION, UNSPECIFIED TISSUE INJURY, OBSTRUCTION/ OCCLUSION, HYPOXIA CANNOT BE DETERMINED. THE REPORTED PATIENT EFFECTS OF TISSUE INJURY AND OBSTRUCTION ARE LISTED IN THE XIENCE SKYPOINT INSTRUCTIONS FOR USE AS KNOWN PATIENT EFFECTS OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE.
20966843 | 2024 | XIENCE XPEDITION 48 EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Unknown | ['Break', 'Material Deformation', 'No Clinical Signs, Symptoms or Conditions'] | MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. IT WAS REPORTED THAT THE PROCEDURE PERFORMED WAS TO TREAT A LEFT ANTERIOR DESCENDING CORONARY ARTERY THAT IS 100% STENOSED. A 2.50X48MM XIENCE XPEDITION WAS IMPLANTED HOWEVER, A STENT STRUT FRACTURE WAS NOTED. A 3.0X33MM XIENCE ALPINE WAS USED TO COVER THE FRACTURED STENT. THERE WERE NO ADVERSE PATIENT EFFECTS AND NO REPORTED CLINICALLY SIGNIFICANT DELAY IN THE PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED. SUBSEQUENT TO THE INITIALLY FILED REPORT, THE FOLLOWING INFORMATION WAS PROVIDED: THE XIENCE XPEDITION IMPLANTED STENT WAS NOTED TO HAVE A SLIGHT BENT STENT STRUT AND NOT FRACTURED. NO ADDITIONAL INFORMATION WAS PROVIDED. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE INCIDENT INFORMATION WAS REVIEWED; HOWEVER, THE PRODUCT WAS NOT RETURNED TO ABBOTT VASCULAR FOR ANALYSIS. FACTORS THAT MAY CONTRIBUTE TO STENT DEFORMATION POST IMPLANTATION INCLUDE, BUT ARE NOT LIMITED TO, STENT RECOIL, IMPLANTATION TECHNIQUE, DAMAGE TO THE STENT, OR VESSEL CHARACTERISTICS. THE PRODUCT RISK ASSESSMENT IDENTIFIES THIS AS A FORESEEABLE EVENT; AS SUCH, DESIGN AND MANUFACTURING CONTROLS AND MITIGATIONS HAVE BEEN ESTABLISHED FOR POTENTIAL CAUSES. BASED ON THE INFORMATION PROVIDED AND BECAUSE THE DEVICE WAS NOT RETURNED FOR ANALYSIS, IT IS UNKNOWN WHAT MAY HAVE CAUSED THE REPORTED STENT DEFORMATION. ANOTHER STENT WAS USED TO COVER THE BENT STENT. A DEFINITIVE CAUSE FOR THE REPORTED ISSUE COULD NOT BE DETERMINED; HOWEVER, THE TREATMENT APPEARS TO BE RELATED TO OPERATIONAL CONTEXT. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. H6 CORRECTION: DEVICE CODE 1069 STENT REMOVED.
20967218 | 2024 | SYNERGY XD | Unknown | ['Material Deformation', 'No Clinical Signs, Symptoms or Conditions'] | IT WAS REPORTED THAT A STENT DAMAGE POST DEPLOYMENT OCCURRED REQUIRING ADDITIONAL INTERVENTION. THE 90% STENOSED TARGET LESION WAS LOCATED IN A MODERATELY TORTUOUS AND MODERATELY CALCIFIED CORONARY ARTERY. THE LESION WAS PRE-DILATED USING A 2X15 EMERGE BALLOON, AND A 2.25 X 38MM SYNERGY XD DRUG-ELUTING STENT WAS ADVANCED AND DEPLOYED. HOWEVER, DURING THE FINAL INTRAVASCULAR ULTRASOUND (IVUS) RUN, THE IVUS CATHETER BECAME CAUGHT ON THE STENT STRUTS. IN RESPONSE, THE PHYSICIAN HAD TO REMOVE THE IVUS CATHETER, BUT IN THE PROCESS, THE STENT BECAME CRIMPED. AS A RESULT, THE PHYSICIAN DEPLOYED A 12MM STENT TO COVER THE DISTAL EDGE OF THE CRIMPED STENT. THE PROCEDURE WAS COMPLETED WITHOUT ANY PATIENT COMPLICATIONS.
20967379 | 2024 | XIENCE XPEDITION 48 EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Unknown | ['Difficult to Advance', 'Vascular Dissection'] | THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED DIFFICULT TO ADVANCE AND SUBSEQUENT TREATMENT APPEAR TO BE RELATED TO OPERATIONAL CONTEXT; HOWEVER, A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT(S), AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THE REPORTED PATIENT EFFECT(S) OF DISSECTION IS LISTED IN THE XIENCE EXPEDITION EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE (EL2098833 REVISION B SECTION 6.2) AS A KNOWN PATIENT EFFECT(S) OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING. IT WAS REPORTED THAT THE PROCEDURE PERFORMED WAS TO TREAT A HEAVILY CALCIFIED, HEAVILY TORTUOUS LEFT CIRCUMFLEX CORONARY ARTERY. WHEN ADVANCING A 2.75X48MM XIENCE EXPEDITION DRUG-ELUTING STENT (DES), RESISTANCE WAS NOTED DUE TO ANATOMY. THE XIENCE EXPEDITION WAS IMPLANTED HOWEVER; A DISSECTION WAS NOTED. AN NEW UNSPECIFIED DES WAS USED TO TREAT THE DISSECTION. NO ADDITIONAL INFORMATION WAS PROVIDED.
20967418 | 2024 | XIENCE SKYPOINTÃ‚Â¿ | Unknown | ['Difficult to Advance', 'Device Dislodged or Dislocated', 'Device Embedded In Tissue or Plaque'] | IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT A LESION IN THE CIRCUMFLEX (CX) ARTERY. THE 2.5X28MM XIENCE SKYPOINT STENT DELIVERY SYSTEM (SDS) HAD RESISTANCE DURING ADVANCEMENT WITH A PREVIOUSLY IMPLANTED STENT IN THE OBTUSE MARGINAL AND A GUIDELINER WAS USED AND REACHED THE TARGET LESION. HOWEVER, THE STENT DISLODGED FROM THE BALLOON AND WAS LEFT IN THE ANATOMY. ANOTHER BALLOON WAS USED TO CRUSH THE DISLODGED STENT AND ANOTHER XIENCE SKYPOINT WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADVERSE PATIENT SEQUELAE AND THERE WAS NO CLINICALLY SIGNIFICANT DELAY IN THE PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. THE INVESTIGATION DETERMINED THE REPORTED DIFFICULTIES AND SUBSEQUENT TREATMENT APPEAR TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE; HOWEVER, A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT(S), AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. IT WAS REPORTED THE STENT DELIVERY SYSTEM (SDS) HAD RESISTANCE DURING ADVANCEMENT WITH A PREVIOUS IMPLANTED STENT RESULTING IN THE REPORTED DIFFICULTY TO ADVANCE. IN THIS CASE, IT IS LIKELY THAT THE MANIPULATION OF THE SDS DURING INTERACTION WITH THE PREVIOUS IMPLANTED STENT CONTRIBUTED TO THE REPORTED DISLODGEMENT OF THE STENT IN THE LESION. ADDITIONALLY, IT WAS REPORTED THAT THE DISLODGED STENT WAS EMBEDDED IN THE PATIENT'S ANATOMY USING A BALLOON CATHETER. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING.
20972272 | 2024 | SYNERGY XD | Unknown | ['Break', 'Inflation Problem', 'Material Deformation', 'Arrhythmia', 'Cardiac Enzyme Elevation', 'Diminished Pulse Pressure'] | IT WAS REPORTED THAT A DEVICE ISSUE AND PATIENT COMPLICATIONS OCCURRED THE 70-80% STENOSED TARGET LESION WAS LOCATED IN THE NON-TORTUOUS AND NON-CALCIFIED CORONARY ARTERY. A 2.50 X 24 MM SYNERGY XD DRUG-ELUTING STENT (DES) WAS ADVANCED EASILY INTO THE VESSEL USING A COPILOT. A SLIGHT BEND WAS NOTICED BUT IT WAS NOTHING OUT OF THE ORDINARY. ALTHOUGH APPROXIMATELY THREE INFLATIONS WERE ATTEMPTED TO REACH NOMINAL PRESSURE, THE DIAL REMAINED AT ZERO AND THE BALLOON DID NOT INFLATE. THE DEVICE WAS REMOVED IN ONE PIECE WITHOUT DIFFICULTY. AFTER THE DEVICE WAS REMOVED, THE PATIENT STARTED HAVING ST ELEVATIONS, VENTRICULAR TACHYCARDIA AND LOST PULSE. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED, THE PATIENT WAS SHOCKED; RESUSCITATION WAS PERFORMED FOR APPROXIMATELY 10 MINUTES TOTAL. THE PATIENT RECOVERED AND THE PROCEDURE WAS COMPLETED SUCCESSFULLY USING AN ALTERNATE DEVICE. THE 2.50 X 24MM SYNERGY WAS INSPECTED AND NOTED TO BE CRACKED IN HALF DOWN THE MIDDLE, ABOUT 1 FOOT FROM THE DISTAL TIP.
20992960 | 2024 | XIENCE XPEDITION 48 EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Unknown | ['Patient-Device Incompatibility', 'Vascular Dissection'] | MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT A 90% STENOSED LESION IN THE PROXIMAL LEFT ANTERIOR DESCENDING (PLAD). THE 3X48MM XIENCE XPEDITION STENT DELIVERY SYSTEM (SDS) WAS ADVANCED TO THE TARGET LESION AND THE STENT WAS IMPLANTED; HOWEVER, A DISSECTION WAS NOTED AND THE STENT WAS ALSO NOTED TO HAVE NOT FULLY EXPANDED AND DID NOT APPOSE ENTIRELY TO THE VESSEL WALL DUE TO LESION CONDITION. THEREFORE, ANOTHER 3X48MM XIENCE XPEDITION STENT WAS IMPLANTED OVER THE INITIALLY IMPLANTED STENT TO TREAT THE DISSECTION. THERE WAS NO ADVERSE PATIENT SEQUELA. NO ADDITIONAL INFORMATION WAS PROVIDED.
20995371 | 2024 | SYNERGY XD | Unknown | ['Difficult to Remove', 'Difficult to Advance', 'Device Dislodged or Dislocated', 'No Clinical Signs, Symptoms or Conditions'] | IT WAS REPORTED THAT STENT DISLODGEMENT OCCURRED. THE 95% STENOSED TARGET LESION WAS LOCATED IN THE LEFT CIRCUMFLEX ARTERY (LCX) WHICH HAD THREE 90-DEGREE BENDS AND A 180-DEGREE ARC OF CALCIFICATION. A 3.00 X 38MM SYNERGY XD DRUG-ELUTING STENT WAS ADVANCED AND ATTEMPTED TO BE DELIVERED THROUGH SIDE CELL OF THE PREVIOUSLY IMPLANTED STENT IN THE LEFT MAIN ARTERY. THE STENT COULD NOT BE DELIVERED FAR ENOUGH DOWN THE LCX SO IT WAS REMOVED WITH THE USE OF AN EXTENSION CATHETER. DURING REMOVAL, THE STENT STRIPPED OFF OF THE BALLOON COMING BACK THROUGH THE SIDE CELL OF THE LEFT MAIN STENT. SECOND ACCESS WAS THEN OBTAINED, AND THE STENT WAS RETRIEVED COMPLETELY USING A SNARE. THE PROCEDURE WAS COMPLETED WITH ANOTHER OF THE SAME DEVICE. NO PATIENT COMPLICATIONS WERE REPORTED.
20995521 | 2024 | SYNERGY MEGATRON | Male | ['Detachment of Device or Device Component', 'Foreign Body In Patient'] | IT WAS REPORTED THAT BALLOON DETACHMENT OCCURRED WHICH REQUIRED INTERVENTION. THE PATIENT PRESENTED WITH ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION. THE TARGET LESION WAS LOCATED IN THE RIGHT CORONARY ARTERY. A 5.00 X 32MM SYNERGY MEGATRON DRUG-ELUTING STENT WAS DEPLOYED FOR TREATMENT. HOWEVER, RIGHT AFTER THE STENT DELIVERY SYSTEM WAS REMOVED, THE PHYSICIAN NOTICED UNDER FLUOROSCOPY THAT THE STENT BALLOON WAS LEFT IN THE PATIENT. SUBSEQUENTLY, MORE STENTS WERE DEPLOYED TO SECURE THE REMAINING STENT BALLOON IN PLACE, AND THE PROCEDURE WAS COMPLETED. NO PATIENT COMPLICATIONS WERE REPORTED.
21006652 | 2024 | ORSIRO MISSION 3.5/15 | Unknown | ['Material Deformation', 'Myocardial Infarction'] | COMBINATION PRODUCT: YES. THE ORSIRO MISSION DRUG-ELUTING STENT WAS IMPLANTED WITHOUT INCIDENCE ON (B)(6) 2024. PATIENT REPRESENTED ON (B)(6) 2024 WITH NSTEMI, AND IN ANGIO DONE ON (B)(6) 2024 STENT IS SEEN TO BE UNDER DEPLOYED IN THE PROXIMAL SECTION. STENT SLIPPED OFF THE BALLOON RESULTING IN UNDER DEPLOYMENT OF THE PROXIMAL SECTION OF THE STENT.
21006662 | 2024 | SYNSIRO PRO 2.5/22 | Unknown | ['Device Dislodged or Dislocated', 'Insufficient Information'] | COMBINATION PRODUCT: YES A SYNSIRO PRO DRUG-ELUTING STENT SYSTEM WAS SELECTED FOR TREATMENT. DURING ATTEMPT TO IMPLANT THE COMPLAINT DEVICE, THE STENT COULD NOT BE POSITIONED. SINCE THE WIRES WERE ENTANGLED, ALL DEVICES WERE REMOVED FROM PATIENTS BODY. AFTER WITHDRAWAL, IT WAS NOTICED UNDER FLUOROSCOPY THAT THE STENT WAS DISLODGED IN THE LEFT MAIN/PROXIMAL LAD. A BYPASS SURGERY WAS PERFORMED. COMBINATION PRODUCT: YES. THE RETURNED PRODUCT WAS SUBJECTED TO A DETAILED TECHNICAL ANALYSIS AND THE CORRESPONDING PRODUCTION DOCUMENTATION WAS REVIEWED TO ESTABLISH WHETHER A DEVIATION FROM THE MANUFACTURING PROCESS COULD BE THE CAUSE FOR THE REPORTED EVENT. IN ADDITION, THE CATHLAB REPORT AND THE ANGIOGRAPHIC MATERIAL PROVIDED WERE REVIEWED. THE TECHNICAL INVESTIGATION CONFIRMED THAT THE STENT HAS DISLODGED FROM THE BALLOON. THE BALLOON IS WELL FOLDED AND SHOWS NO SIGNS OF INFLATION. STENT IMPRINTS WERE OBSERVED ON THE EXPOSED BALLOON SURFACE, INDICATING THAT THE STENT WAS CRIMPED IN BETWEEN THE TWO RADIOPAQUE MARKERS AT THE TIME OF DELIVERY. THE STENT WAS NOT RETURNED. THE ANGIOGRAPHIC MATERIAL SHOWS THE TREATMENT TARGET LESION. THE DISLODGED STENT REMAINS IN THE DISTAL LM/PROXIMAL LAD. THE ACTUAL COMPLAINT EVENT IS NOT VISIBLE. THE ANGIOGRAPHIC MATERIAL DOES THEREFORE NOT PROVIDE FURTHER RELEVANT INFORMATION WITH REGARDS TO THE ROOT CAUSE OF THE REPORTED EVENT. REVIEW OF THE PRODUCTION DOCUMENTATION CONFIRMED THAT THE PRODUCT WAS MANUFACTURED ACCORDING TO SPECIFICATIONS AND PASSED ALL IN-PROCESS AND FINAL INSPECTIONS. DURING FINAL INSPECTION, EVERY STENT SYSTEM UNDERGOES VISUAL INSPECTION TO ENSURE CORRECT STENT EMBEDDING AND HOMOGENEOUS CRIMPING OF THE STENT. FURTHER, THE STENT OUTER DIAMETER OF A DEFINED AMOUNT OF SAMPLES IS TESTED FROM EVERY LOT AND THE STENT RETENTION FORCE IS TESTED BY MEANS OF MANUFACTURING PROCESS OUTPUT MONITORING. BASED ON THE CONDUCTED INVESTIGATIONS OF THE DEVICE BEING SUBJECT TO THIS COMPLAINT, NO MATERIAL OR MANUFACTURING RELATED ROOT CAUSE COULD BE DETERMINED.

## Sheet2
eval_data_pretrained_llama_results
Evaluation metrics: {'rougeL': 0.5621860040688778, 'exact_match': 0.3869209809264305, 'custom_match': 0.3869209809264305, 'f1_score': 0.712451532361293, 'accuracy': 0.44686648501362397}
eval_data_finetuned_llama_results
Evaluation metrics: {'rougeL': 0.8594608209170556, 'exact_match': 0.6811989100817438, 'custom_match': 0.6811989100817438, 'f1_score': 0.9145850343502835, 'accuracy': 0.771117166212534}
eval_data_finetuned_GPT2_results
Evaluation metrics: {'rougeL': 0.9891593481256571, 'exact_match': 0.9455040871934605, 'custom_match': 0.9455040871934605, 'f1_score': 0.9900862650910255, 'accuracy': 0.9754768392370572}
eval_data_pretrained_gpt2_results
Evaluation metrics: {'rougeL': 0.02664686652891008, 'exact_match': 0.01634877384196185, 'custom_match': 0.01634877384196185, 'f1_score': 0.35021015965158997, 'accuracy': 0.027247956403269755}
| Model | BERT-score
| Llama (Base) + RAG | 0.71245
| Llama (Fine-tuned) + RAG | 0.91459
| GPT2 (Base) + RAG | 0.35021
| GPT2 (Fine-tuned) + RAG | 0.99009